## Whitepaper
Whitepaper: BOT-3000E Tribometer and Compliance to ASTM F2508-23 Standard
Introduction
The BOT-3000E tribometer, engineered and manufactured by Regan Scientific / Walkway Management Group, has established itself as the leading device for accurately measuring the Dynamic Coefficient of Friction (DCOF) on hard surface flooring. Its adoption into ANSI standards since 2012 (ANSI A137.1) and subsequently ANSI A326.3 in 2017 underscores its critical role in floor safety standards and practices. With recent inclusion into the 2024 International Swimming Pool and Spa Code (ISPSC), the BOT-3000E continues to demonstrate unparalleled relevance in slip resistance assessment.
ASTM F2508-23 Compliance
Walkway Management Group proudly announces the compliance of the BOT-3000E with the latest ASTM F2508-23 standard.
The F2508-23 standard outlines a rigorous validation process based on human-subject ambulation study, requiring tribometers to meet stringent criteria, including reproducibility, repeatability, and precise surface friction ranking. To achieve this, an Interlaboratory Study (ILS) involving nine independent laboratories assessed the BOT-3000E equipped with a newly developed sensor. This sensor has been developed strictly for the F2508 standard and should not be used for any ANSI audits.
Over thousands of measurements across 12 distinct sets of tiles—each set comprising three pre-selected tiles—BOT-3000E consistently delivered flawless ranking aligned with human-subject trials. Impressively, not a single overlapping measurement between tile rankings occurred, demonstrating absolute precision.
Compliance to ANSI A326.3 and ISPSC 2024 / USPSHTC 2024
The versatility of the BOT-3000E is reinforced by its alignment with ANSI A326.3, recognized as the most scientifically robust standard for DCOF measurement. ANSI A326.3 provides detailed slip resistance thresholds and informative values tailored specifically for various flooring applications, resulting in safer, evidence-based floor selection and maintenance practices. The integration of ANSI A326.3 within the ISPSC 2024 further and USPSHTC 2024 validates the BOT-3000E’s critical role in promoting pool and spa safety, reflecting a broader international acceptance and trust in its accuracy and reliability.
ANSI A326.3: The Gold Standard
Since its 2017 release, ANSI A326.3 revolutionized the approach to flooring safety by introducing scientifically-backed, application-specific DCOF thresholds and informative values . Its significance is underscored by decades of thorough research and validation studies, positioning ANSI A326.3 as the gold standard in slip resistance measurement.
Though compliance with ASTM F2508-23 remains valuable in demonstrating tribometer consistency and precision, the scientific depth and practical applicability of ANSI A326.3 firmly establish it as the authoritative guide for slip resistance standards in the United States and internationally.
BOT-3000E and the Daubert Principles
Critical to litigation and forensic scenarios, the BOT-3000E satisfies the Daubert Principles by providing scientifically robust, reproducible, and widely accepted measurements of slip resistance. Its automated and verifiable testing process significantly enhances its reliability and evidentiary credibility in court. Moreover, the automated reports can be verified for tampering. Another great feature of the device is that it ensures that the data presented is 100% accurate and non-altered.
Device Versatility and Technological Excellence
Beyond its technical compliance, the BOT-3000E excels through unparalleled portability, user-friendly operations, and automated report generation capabilities, which simplify slip resistance assessments across various environments. Its intuitive design ensures consistent and precise operation, reaffirming its position as the most versatile and technologically advanced tribometer available today.
Conclusion
The BOT-3000E continues to lead the field in slip resistance measurement, validated through its recent compliance with ASTM F2508-23, ANSI A326.3, and inclusion in the ISPSC 2024 / USPSHTC 2024. Its scientifically rigorous design and operational ease underscore its invaluable role in ensuring walkway safety and reducing slip-and-fall incidents globally.
Walkway Management Group remains committed to innovation, precision, and safety, ensuring that the BOT-3000E remains the preferred tribometer for professionals dedicated to walkway safety.

## F2508 Tile Results
RESULTS
The data shared in the previous tabs and below has been acquired following Standard ASTM F2508-23 and results have been calculated using Standard ASTM E691. The BOT-3000E in conjunction with the F2508 sensor/slider complies with all requirements of F2508-23 for Rank Order, Differentiation as well as Established limits as defined in section 9.1.1, 9.1.2 and 9.1.3 of the standard.
Tile E |  | Tile F |  | Tile G
xdbar | 0.079 | xdbar | 0.226 | xdbar | 0.597
sr | 0.01046 | sr | 0.0077028 | sr | 0.01491
r | 0.029287 | r | 0.021568 | r | 0.041747
sR | 0.017152 | sR | 0.022855 | sR | 0.066223
R | 0.048026 | R | 0.063993 | R | 0.18542
R* | 0.035437 | R* | 0.047218 | R* | 0.13682
low limit | 0.043591 | low limit | 0.17896 | low limit | 0.46038
high limit | 0.11446 | high limit | 0.2734 | high limit | 0.73401
number of TR below xdbar-R*: |  | number of TR below xdbar-R*: |  | number of TR below xdbar-R*:
1 |  | 1 |  | 2
number of TR above xdbar+R*: |  | number of TR above xdbar+R*: |  | number of TR above xdbar+R*:
4 |  | 2 |  | 0
* Tests were conducted using WMG's "F2508 sensor", and the data presented is not associated with the ANSI A326.3 standard. For statistical analyses pertaining to the ANSI A326.3 standard, please refer to Table 2 in the latest version of the standard.

## F2508 ILS Form
INTRODUCTION
This document is meant to be a test procedure designed for the laboratories of the Super ILS hosted in Baltimore in April 2024. As this is a quick reference for the users of each laboratory, not all elements of the F3539 have been included, for clarity and ease of use. A more comprehensive Test Procedure, following ASTM F3539 requirements will be available after the ILS is completed.
Scope of testing
I order to complete the ILS, each Lab (consisting of a user, tribometer and a slider) will need to provide testing data on the following:
· 12 sets of Reference Materials (RM)
· Each set consists of 3 different tiles (E, F and G)
· Each tile will require 3 test results
· Each Test result will consist of the average of 4 observations (4 single BOT run of 8”), ran in 4 different directions on the RM (E, W, N and S).
In total, this will mean 10 x 3 x 3 x 4 = 360 individual observations/runs.
The slider will need to be reconditioned after each test result is completed (every 4 observations). The main purpose of reconditioning is to ensure that no wear line is showing on the slider before starting the process for the next test result.
The sandpaper shall be changed after every 3 RMs
Items provided by the ILS organizers
· Clean tiles
· SLS
· Microfiber towels
· Lab assistant/observer
Items provided by WMG for its Labs
· F2508 sensors
· Sandpaper specific to the F2508 sensors (10 each)
· Plastic risers (12” x 4”) to put on each side of the tile that will be under test allowing the BOT to run its 8” course. Note, each participant will have 2 sets of risers as there are two different thickness of tiles to be tested. The thicker version (various color) will be used for tile E and G, the thinner version (gray) will be used on the thinner tile F. It will be obvious as of what risers go with what tile as the height of the risers and the RM must be even.
· Data recording forms
Data Recording
Each lab will receive a form for data completion. The document is structured in a way where it will be easy to capture all the required data. Make sure you enter the data clearly and return your forms to the coordinator at the end of the ILS.

## F3539 ILS Validation
No. | Description | Notes
A1.1 | ILS General Information
A1.1.1 | Date or date range of ILS. | 2024-04-28
A1.1.2 | Operator name and contact information for each laboratory. | Yves Lafortune, y.lafortune@walkwaymg.com Shandra Winchester, s.winchester@walkwaymg.com Ian Kenny, i.kenny@walkwaymg.com Antonia Welker, a.wekler@walkwaymg.com Matthew Idler, midler@mooreeng.com Manny Winchester, m.winchester@walkwaymg.com Michael Reighn, mreighn@mooreeng.com Grand Davidson, gdavidson@tileusa.com Scott Davis,sdavis@ipalaboratories.com Brian Vassallo, b.vassallo@safespaceingenuity.com
A1.1.3 | ILS ambient temperature/humidity range. | 71F, 60% RH
A1.1.4 | Serial numbers of tribometers used in ILS. | 121, 435, 57, 560, 569, 387, 259, 258, 442
A1.1.5 | Serial numbers of sliders used in ILS. | 865, 864, 821, 820, 913, 191, 832, 855, 831
A1.1.6 | Number of valid lab data sets (see Practice E691 Sections 17 - 19). | 9 Labs
A1.2 | Tribometer Information
A1.2.1 | Tribometer supplier: | Walkway Management Group
A1.2.1.1 | Mailing address. | 400 Gerault Rd, Flower Mound TX 75028
A1.2.1.2 | Phone numbers. | 1-866-964-9255
A1.2.1.3 | Website. | www.walkwaymg.com
A1.2.2 | Tribometer model (unique): | BOT-3000E
A1.2.2.1 | Slider (complete) part identifier. | F2508 Sensor
A1.2.2.2 | Serial numbering method for each slider. | S/N in the slider/sensor that can be retrieved through the menu.
A1.2.3 | Tribometer energy source consumables – description: | The BOT-3000E uses a NmH (Nickle Metal Hydride) battery.
A1.2.3.1 | Supplier and website. | Walkway Management Group - www.walkwaymg.com
A1.2.3.2 | Supplier part identifier. | 9016212
A1.2.4 | Tribometer energy source setting range. | N/A
A1.2.5 | Slider frictional material – description: | SBR 75A
A1.2.5.1 | Supplier and website. | Walkway Management Group - www.walkwaymg.com
A1.2.5.2 | Supplier part identifier. | Not available for standalone purchase. Please refer to complete slider for purchase.
A1.2.5.3 | Overall dimensions of frictional material (mm). | 28 x 28
A1.2.5.4 | Contact surface feature (for example, grooving) dimensions. | N/A
A1.2.6 | Slider resurfacing tools (not consumables) – description: | Hand-operated sanding tool for reconditioning.
A1.2.6.1 | Supplier and website. | Walkway Management Group - www.walkwaymg.com | •
A1.2.6.2 | Supplier part identifier. | 9013057
A1.2.7 | Slider resurfacing consumables – description: | Sandpaper Black Carbon - 1,000 Grit
A1.2.7.1 | Supplier and website. | Walkway Management Group - www.walkwaymg.com
A1.2.7.2 | Supplier part identifier. | 9020171
A1.2.8 | Slider cleaning materials – description: | Brush
A1.2.8.1 | Supplier and website. | Walkway Management Group - www.walkwaymg.com
A1.2.8.2 | Supplier part identifier. | 2019565
A1.2.9 | Instructions for actuating tribometer to obtain a test observation. | 1. Insert a fully charged battery pack into the battery compartment located at the front of the device. The latch should engage to secure the battery pack to the device. 2. Tip the device onto either side or back panel and fully insert the desired sensor (slider) into the receiver located in the bottom of the device. Return the device to its operating position (wheels down). 3. Center the device directly over the test area, noting the approximate location of the sensor (side arrow labels). When performing wet tests, be certain the sensor touches down onto the applied liquid film. Avoid getting the wheels wet, which could induce a wheel slip condition. Press any switch momentarily to power up the device. There will be a slight delay before the display illuminates. 4. Press the DCOF switch. 5. In DCOF mode, you select “SINGLE” to perform one test. 6. Use the navigation switches (arrows, + and -) to adjust the test run distance to 8". 7. Press the “GO” switch to begin testing. The unit will emit a beep alert, lower the sensor, and then move forward. CAUTION: To protect the internal mechanism, do not lift or move the instrument from the test surface until the beep alert stops! 8. The device will show the DCOF measured and also save it in its memory.
A1.3 | ILS Reference Material Information (for each material used) | USC3 Tile E
A1.3.1 | Reference material description: | Porcelain tile with polished finish
A1.3.1.1 | Material and finish (bare substrate). | American Wonder Porcelain, Orvieto, color OR01 Ivory, 305mm (12in.) square
A1.3.1.2 | Supplier and website. | ASTM International, ASTM.org
A1.3.1.3 | Supplier part identifier. | ADJF2508AWP
A1.3.1.4 | Date obtained. | 2023-08-01
A1.3.1.5 | Quantity used in ILS. | 12
A1.3.1.6 | Lot/batch/serial number for each material used in ILS. | 55504
A1.3.1.7 | One of the following statements: (1) A citation to specific documentation of the homogeneity and stability of the reference material. (2) A statement that the homogeneity and stability of the reference material has not been documented. | Homogeneity of the reference material has not been provided to Walkway Management Group as of the time this report has been written.
A1.3.1.8 | Reference material supplemental coating/finish: (1) Supplier and website. (2) Supplier part identifier. (3) Application instructions used for ILS materials. | N/A
A1.3 | ILS Reference Material Information (for each material used) | USC3 Tile F
A1.3.1 | Reference material description: | Ceramic tile with glazed finish
A1.3.1.1 | Material and finish (bare substrate). | Daltile, Sierra, color 1500 Rainier, 305mm (12in.) square
A1.3.1.2 | Supplier and website. | ASTM International, ASTM.org
A1.3.1.3 | Supplier part identifier. | ADJF2508DAT
A1.3.1.4 | Date obtained. | 2023-08-01
A1.3.1.5 | Quantity used in ILS. | 12
A1.3.1.6 | Lot/batch/serial number for each material used in ILS. | 16008
A1.3.1.7 | One of the following statements: (1) A citation to specific documentation of the homogeneity and stability of the reference material. (2) A statement that the homogeneity and stability of the reference material has not been documented. | Homogeneity of the reference material has not been provided to Walkway Management Group as of the time this report has been written.
A1.3.1.8 | Reference material supplemental coating/finish: (1) Supplier and website. (2) Supplier part identifier. (3) Application instructions used for ILS materials. | N/A
A1.3 | ILS Reference Material Information (for each material used) | USC3 Tile G
A1.3.1 | Reference material description: | Porcelain tile with matte finish
A1.3.1.1 | Material and finish (bare substrate). | Crossville, Virtue, color white, 305mm (12in.) square
A1.3.1.2 | Supplier and website. | ASTM International, ASTM.org
A1.3.1.3 | Supplier part identifier. | ADJF2508CRO
A1.3.1.4 | Date obtained. | 2023-08-01
A1.3.1.5 | Quantity used in ILS. | 12
A1.3.1.6 | Lot/batch/serial number for each material used in ILS. | AV261.mc…3588
A1.3.1.7 | One of the following statements: (1) A citation to specific documentation of the homogeneity and stability of the reference material. (2) A statement that the homogeneity and stability of the reference material has not been documented. | Homogeneity of the reference material has not been provided to Walkway Management Group as of the time this report has been written.
A1.3.1.8 | Reference material supplemental coating/finish: (1) Supplier and website. (2) Supplier part identifier. (3) Application instructions used for ILS materials. | N/A
A1.4 | Cleaning Supplies for ILS Reference Materials (for each cleaner used)
A1.4.1 | Cleaner solute (or solution if sold premixed): | Sodium lauryl sulfate (SLS) liquid, 30% pure
A1.4.1.1 | Supplier and website. | Chemistry Store, www.chemistrystore.com
A1.4.1.2 | Supplier part identifier. | 82025 (1 gallon)
A1.4.2 | Cleaner solvent: | Distilled water
A1.4.2.1 | Supplier and website. | Brand nonspecific
A1.4.2.2 | Supplier part identifier. | Brand nonspecific
A1.4.2.3 | Mixing proportions. | 6.3mL SLS in 3.785L (1gal) distilled water = 0.05% SLS solution
A1.4.3 | Cleaner application tools: | Microfiber towels
A1.4.3.1 | Supplier and website. | Brand nonspecific
A1.4.3.2 | Supplier part identifier. | Brand nonspecific
A1.5 | Contaminant Information (for each contaminant used)
A1.5.1 | Contaminant solute name: | Sodium lauryl sulfate (SLS) liquid, 30% pure
A1.5.1.1 | Supplier and website. | Chemistry Store, www.chemistrystore.com
A1.5.1.2 | Supplier part identifier. | 82025 (1 gallon)
A1.5.2 | Contaminant solvent name: | Distilled water
A1.5.2.1 | Supplier and website. | Brand nonspecific
A1.5.2.2 | Supplier part identifier. | Brand nonspecific
A1.5.2.3 | Mixing proportions. | 6.3mL SLS in 3.785L (1gal) distilled water = 0.05% SLS solution
A1.5.3 | Contaminant application tools: | Bottle flip-top screw on lid - 8oz
A1.5.3.1 | Supplier and website. | Walkway Management Group - www.walkwaymg.com
A1.5.3.2 | Supplier part identifier. | 2023082
A1.6 | Friction Test Procedure for ILS
A1.6.1 | Reference material cleaning method. | Ref ASTM F2508-2023 section 8.3.1.1 - 8.3.1.7: 8.3.1.1 No surface treatment except as specified in this section is permitted. 8.3.1.2 Disperse at least 25 mL of 0.05 % SLS solution onto the RM surface. Gently wipe with a clean microfiber cloth, in a circular pattern, the entire RM test surface for a minimum of 10-s. Reapply 25 mL of 0.05 % SLS and repeat the minimum 10-s of wiping with the same microfiber cloth. Ensure that the 0.05 % SLS solution has been applied to the entire RM test surface. NOTE 1—To clean the microfiber cloth after use, run a basin of cool or warm water, insert the cloth and agitate it by hand. Soak the cloth for a minimum of 15 min then agitate it again by hand. Remove it from the water basin and thoroughly rinse under cool or warm running water. Wring out as much remaining liquid as possible by hand, then hang dry. 8.3.1.3 Rinse the surface thoroughly with distilled or deionized water, ensuring that no visible suds or residues remain. 8.3.1.4 Prepare an ethanol solution containing equal parts denatured ethanol and distilled water. 8.3.1.5 Disperse at least 25 mL of ethanol solution onto the RM surface. With a clean microfiber cloth different from that used in 8.3.1.2, wipe the RM test surface for 10-s. 8.3.1.6 Dry the RM with dry and oil-free compressed air or air dry if compressed air is not available. Any visible contamination remaining after this step will disqualify the RM for use. 8.3.1.7 Ensure that handling of the RM does not introduce contaminants to the test surfaces, including exposing the test surfaces to contact by human skin.
A1.6.2 | Reference material cleaning frequency. | Each reference material unit is to be cleaned before each lab's test.
A1.6.3 | Method for slider cleaning. | - Before reconditioning, a slider should be dried on a dry microfiber to keep the sandpaper free of contaminant. - Before the '[Contact_Hours_Vario_2024.xlsx]4'!$C$133first test observation of a test result group is started, the slider should be cleaned with a damp microfiber.
A1.6.4 | Slider resurfacing method (for example, sanding). | 1. Use fresh sandpaper every 3 test results. Thoroughly brush the sandpaper to remove any loose sand particles from the paper. 2. Insert the sensor and sand until any noticeable wear pattern or flat spot has been removed. 3. Remove sensor and wipe off debris with a clean, lint free, oil free paper towel or clean brush. 4. Thoroughly brush all the debris from the sand paper.
A1.6.5 | Slider resurfacing frequency (for example, number of slips). | Every 4 observations.
A1.6.6 | Slider post-resurfacing/pre-testing “break-in” procedure. | Clean per item A1.6.3.
A1.6.7 | Contaminant application instructions. | Using a squeeze bottle and a small clean paint brush, apply a continuous film covering the length of the run of the appropriate contaminant, per item A1.5, completely down the center of the reference surface approximately 3 inches wide.
A1.6.8 | Test position(s) on reference tiles (with dimensions if not centered). | The observation should be centered on the RM. Depending on the size of the RM, it may require risers, like other RM, on each side of the RM being tested to ensure that the BOT-3000E's wheel have an even plane throughout the complete the observation.
A1.6.9 | Test direction(s) on reference tiles (initial and rotation degrees). | The 4 observations should be completed with the following order for best efficiency, according to the markings on the RM: 0, 180, 90 and 270.
A1.6.10 | Number of test observations in each direction. | 1 observation in each direction.
A1.7 | Data Conversion Instructions
A1.7.1 | Instructions for converting a test observation (or observations) into a test result. | The average of 4 test observations converts in a test result.
A1.7.2 | Instructions for converting measurements obtained on sloped surfaces. | N/A

## F3539 Field Testing
No. | Description | Notes
A1.8 | General Information for Field Testing
A1.8.1 | Tribometer serial number range that test procedure applies to. | All BOT-3000E S/N
A1.8.2 | Tribometer storage requirements if not indoor climate controlled. | 5 - 40 deg. C
A1.8.3 | Allowable temperature and humidity ranges for use. | 5 - 40 deg. C , 5% - 90% RH (non-condensing)
A1.8.4 | Tribometer maintenance required of user. | The BOT-3000E requires a minimum amount of maintenance as a function of design, however there are several items that should be noted: BATTERIES: - Battery packs should never be transported or stored where metallic objects or liquids may contact the connection interface or it could arc and ignite any surrounding material. - Leaving a battery pack inside the device or the charger when not in use will drain the pack to an unusable state. - Battery packs should be maintained on a rotating charging schedule to prevent premature failure. - Battery pack charging should take place in an open area free from excessive heat, dust, and flammable gasses or liquids. - Batteries may become quite warm during charging, and must be allowed to cool completely before using.
A1.8.5 | Recommendations for and frequency of laboratory verification. | Yearly calibration by the manufacturer required
A1.8.6 | Recommendations for and frequency of user verification. | The Verification operation is intended to determine if the internal force measurement system is within specification. It is NOT a calibration procedure. This must be done when getting on-site, before testing a surface. 1. Allow the tribometer to acclimate to the ambient temperature of the test area. 2. Press any one of the six panel switches momentarily to power up the device. 3. Place the device onto a very stable and level surface. 4. Stand the device on the aft (printer) end, and then carefully insert the verification mass into the sensor port. 5. The VERIFY option indicator will appear (Fig.2, lower left screen). 6. Press the switch adjacent to “VERIFY”. NOTE: Near the screen bottom, there are two tilt indicators. These should indicate 90 +/- 2 degrees and 0 +/- 2 degrees in order for the device to verify accurately, otherwise a more level surface must be found. 7. Press the “START” switch; the device will now begin the verification process. Be certain that the device remains level and stable during this process, otherwise the device may fail. Once the device has completed the verification process, it will indicate a “PASS” or a “FAIL” status. You may press the “PRINT” or “DONE” switch at this time. If the device fails, the procedure may be repeated after rechecking the level or moving to a more level location. The mass may require removing and replacing or be checked for free movement. If the device passes verification, the force measurement system is within factory specifications. 8. Wait several seconds before removing the verification mass so the device can reset. 9. If the device cannot be verified because of an internal fault, it must be returned to Walkway Management Group. The F2508 Sensor Validation Procedure is recommended to be completed before and after every audit to ensure proper maintenance of the F2508 Sensor. 1. The ASTM F2508-2023 adjunct reference material Tile G can be used for the verification step. 2. The tile shall be clean and free of contamination prior to use. 3. Clean the tile using denatured alcohol or light-use cleaner, like Bona. 3. If cleaning with denatured alcohol, disperse onto the tile surface and wipe using a clean, dry microfiber cloth or paper towel that does not leave residue to remove any dust, debris, or other contaminants that have collected on the surface. Rinse thoroughly with water to remove any remaining material. 4. If cleaning with light-use cleaner, disperse onto the tile surface and wipe using a clean, dry microfiber cloth or paper towel that does not leave residue to remove any dust, debris, or other contaminants that have collected on the surface. Rinse thoroughly with water to remove any remaining material. 5. Allow to air dry or dry with a clean, dry microfiber cloth or paper towel that does not leave residue. 6. Wet the path that the slider will follow with enough 0.05% SLS water such that the slider will remain within the wetted path throughout the entire test. 7. With a resurfaced sensor/slider, make one 8 in. (203.2 mm) measurement then rotate the testing device 180° and make a second 8 in. (203.2 mm) measurement on the validation surface. Add more SLS if needed between passes. Repeat a 2nd time for a total of 4 measurements. Calculate the average. 8. If the average is within the low and high limits of Tile G as specified by the results of the ILS as found in this document, the F2508 sensor validation is deemed validated for the audit. If the value falls above or below that range, repeat the test up to 3 times. Repeated failure may be require a new sensor/slider.
A1.8.7 | Information on the useful life of apparatus: | Varies by utilization frequency and environment. Determined by calibration/verification issues. Over 10 years, in average.
A1.8.7.1 | Slider useful life. | Over 2.5mm thickness. No sign of wear and tear on the surface after re-conditioning.
A1.8.7.2 | Contaminant(s) useful life. | SLS is acceptable to use as long as the solution is uniform, so separation. Keep at room temperature in dark area.
A1.8.7.3 | Cleaner useful life. | Same as A1.8.7.2
A1.8.8 | Maximum slope allowable for friction testing. | 9.6 degrees Maximum
A1.9 | Reference Material Information for Use in User Verification (for each material used)
A1.9.1 | Reference material description: | ASTM F2508-2023 adjunct reference material Tile G
A1.9.1.1 | Material and finish (bare substrate): (1) Supplier and website. (2) Supplier part identifier. | Porcelain floor tile, matte finish Full Tiles: ASTM (www.astm.org) ADJF2508CRO Strips: Walkway Management Group (www.walkwaymg.com) F2508 - Tile Bundle (Includes E, F, G)
A1.9.1.2 | Method of unique identification. | N/A
A1.9.1.3 | Useful life duration if specified. | N/A
A1.9.1.4 | Reference material supplemental coating/finish: (1) Supplier and website. (2) Supplier part identifier. (3) Application instructions. | N/A
A1.9.2 | Reference material cleaning method.
A1.9.3 | Reference material cleaning frequency. | Before every use.
A1.9.4 | Test position(s) on reference tile (with dimensions if not centered). | Centered
A1.9.5 | Test direction(s) on reference tile (initial and rotation degrees). | E, W (full tiles must be marked for orientation)
A1.9.6 | Number of tests in each direction. | 2
A1.10 | Cleaning Supplies for User Verification Reference Material(s) | Microfiber cloth or paper towel: shall not leave lint or residue after use.
A1.10.1 | Cleaner solute (or solution if sold premixed): | Light-use cleaner: Shall be commercially available, pH neutral, water-based, light-use, non-degreasing, non-disinfecting, and non-wax stripping. Cleaner shall not leave surface residue after wiping with a microfiber cloth or paper towel nor alter the test specimen surface. Bona® Stone, Tile and Laminate Cleaner or the equivalent has been found satisfactory. It is commercially available at various retail locations. Denatured alcohol.
A1.10.1.1 | Supplier and website. | Bona Stone, Tile & Laminate Floor - Pro Series (www.bona.com)
A1.10.1.2 | Supplier part identifier. | WM700051188
A1.10.2 | Cleaner solvent: | N/A
A1.10.2.1 | Supplier and website. | N/A
A1.10.2.2 | Supplier part identifier. | N/A
A1.10.2.3 | Mixing proportions. | N/A
A1.10.3 | Cleaner application tools: | Generic Microfiber cloth
A1.10.3.1 | Supplier and website. | N/A
A1.10.3.2 | Supplier part identifier. | N/A
A1.10.4 | Useful life duration from new & storage requirements. | 2 year shelf life for Bona. N/A for denatures alcohol.
A1.11 | Friction Test Procedure for User Verification
A1.11.1 | Method for slider cleaning. | Slider can be wiped on a clean microfiber rag to eliminate any sanding residue.
A1.11.2 | Slider resurfacing method (for example, sanding). | Sanding. Follow sanding process per manufacturer's instructional video in FAQ. https://www.walkwaymg.com/faqs/?ver=1.0. Sandpaper Black Carbon - 1,000 Grit is required for the F2508 slider conditioning. Can be purchased on manufacturer's webstore.
A1.11.3 | Slider post-resurfacing/pre-testing “break-in” procedure. | N/A
A1.11.4 | Contaminant application instructions. | SLS should be distributed on the tile strip or along the slider path on a full size tile with a dispensing bottle. Soft bristle brush should be used to ensure complete coverage of the tile with SLS.
A1.11.5 | Test position(s) on reference tiles (with dimensions if not centered). | Centered
A1.11.6 | Test direction(s) on reference tiles (initial and rotation degrees). | E, W (full tiles must be marked for orientation)
A1.11.7 | Number of test observations in each direction. | 2 for the strip, 1 for the large tile. Total of 4 observations.
A1.11.8 | Slider resurfacing frequency (for example, number of slips). | After 4 slips
A1.12 | Friction Test Procedure for Field Testing
A1.12.1 | Instructions for selecting testing locations. | Auditor will need to select 3 adjacent locations of the walkway surface under test. These locations should represent different wear patterns. If the test is completed under prevalent conditions, different level of contamination should be measured.
A1.12.2 | Instructions for determining if a surface is inappropriate for testing. | Three-dimensionally patterned or profiled surfaces can affect device readings as well as device mobility. Size and shape of the patterns have a direct effect on the ability of the device to complete a run. An incomplete run will be noted on the display with an error message. During a successful run on profiled surfaces, the auditor must take into considerations the effect of patterns for their assessment of the slip resistance. Grout lines of tiles should be ran over at a 45 degree angle by the tribometer, when possible, to avoid possible wheel spin.
A1.12.3 | Instructions for deciding on test directions and number of tests in each direction. | When space allows, 4 runs (N,S, E, W) must be completed and averaged for the 1st test sample result. 2 more test samples must be completed per A1.12.1. The lowest value of the 3 test samples shall be selected for the final result of the audit. Smaller areas, like stairs, can be tested by completing the 4 runs by repeating the East, West runs twice.
A1.12.4 | Instructions for verifying a slider frictional surface is usable. | Ensure that the rubber is more than 2.5mm and no visible wear line after resurfacing. Rubber must be completely secured to the slider.
A1.12.5 | Method for testing on slopes. | Slopes should not exceed that maximum grade mentioned in A1.8.8 and be tested N, S, E, W.
A1.12.6 | Method for slider cleaning. | Slider can be wiped on a clean microfiber rag to eliminate any sanding residue.
A1.12.7 | Slider resurfacing method (for example, sanding). | Sanding. Follow sanding process per manufacturer's instructional video in FAQ. https://www.walkwaymg.com/faqs/?ver=1.0. Sandpaper Black Carbon - 1,000 Grit is required for the F2508 slider conditioning. Can be purchased on manufacturer's webstore.
A1.12.8 | Slider resurfacing frequency (for example, number of slips). | After 4 slips
A1.12.9 | Slider post-resurfacing/pre-testing “break-in” procedure. | N/A
A1.12.10 | Walkway surface cleaning requirements. | If surface requires cleaning, the auditor should select the proper chemical that will remove the contaminants present. The cleaning process should be concluded with a neutral cleaner wash like the one used for the validation surface and a final water rinse, to eliminate any cleaner residue.
A1.12.11 | Contaminant application instructions. | SLS should be distributed on the slider path of a test run with a dispensing bottle. Soft bristle brush should be used to ensure complete coverage of the sensor trajectory with SLS.

## F2508 Data
|  |  | Test 1 |  |  |  | Test 2 |  |  |  | Test 3 |  |  |  | Averages
Lab | Seq | ID | East | West | North | South | East | West | North | South | East | West | North | South | Test 1 | Test 2 | Test 3 | Notes
1 | 16 | E01 | 0.08 | 0.11 | 0.06 | 0.07 | 0.07 | 0.1 | 0.06 | 0.07 | 0.09 | 0.12 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09
2 | 11 | E01 | 0.06 | 0.06 | 0.08 | 0.07 | 0.05 | 0.05 | 0.08 | 0.06 | 0.06 | 0.06 | 0.08 | 0.06 | 0.07 | 0.06 | 0.07
3 | 31 | E01 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.08 | 0.06 | 0.07 | 0.08 | 0.07 | 0.06 | 0.06 | 0.08 | 0.07 | 0.07
4 | 28 | E01 | 0.11 | 0.08 | 0.06 | 0.06 | 0.12 | 0.08 | 0.06 | 0.06 | 0.07 | 0.07 | 0.06 | 0.06 | 0.08 | 0.09 | 0.07
5 | 4 | E01 | 0.1 | 0.08 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.08 | 0.06 | 0.05
6 | 35 | E01 | 0.09 | 0.09 | 0.06 | 0.06 | 0.08 | 0.08 | 0.05 | 0.06 | 0.08 | 0.08 | 0.06 | 0.06 | 0.08 | 0.07 | 0.07
7 | 26 | E01 | 0.06 | 0.07 | 0.06 | 0.08 | 0.07 | 0.07 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.07 | 0.07 | 0.05
8 | 24 | E01 | 0.08 | 0.12 | 0.07 | 0.07 | 0.08 | 0.09 | 0.08 | 0.14 | 0.08 | 0.09 | 0.08 | 0.08 | 0.09 | 0.08 | 0.08
9 | 36 | E01 | 0.1 | 0.08 | 0.11 | 0.06 | 0.13 | 0.09 | 0.1 | 0.11 | 0.15 | 0.08 | 0.09 | 0.08 | 0.09 | 0.11 | 0.1
1 | 15 | E02 | 0.11 | 0.11 | 0.09 | 0.09 | 0.12 | 0.13 | 0.1 | 0.11 | 0.1 | 0.13 | 0.09 | 0.1 | 0.1 | 0.12 | 0.11
2 | 33 | E02 | 0.07 | 0.06 | 0.12 | 0.05 | 0.09 | 0.05 | 0.13 | 0.06 | 0.13 | 0.06 | 0.13 | 0.06 | 0.08 | 0.09 | 0.1
3 | 34 | E02 | 0.1 | 0.09 | 0.09 | 0.09 | 0.09 | 0.08 | 0.09 | 0.07 | 0.08 | 0.07 | 0.08 | 0.07 | 0.09 | 0.09 | 0.08
4 | 32 | E02 | 0.07 | 0.06 | 0.05 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.1 | 0.09 | 0.06 | 0.06 | 0.06 | 0.06 | 0.08
5 | 1 | E02 | 0.07 | 0.06 | 0.07 | 0.06 | 0.06 | 0.06 | 0.07 | 0.06 | 0.05 | 0.05 | 0.06 | 0.06 | 0.07 | 0.06 | 0.06
6 | 14 | E02 | 0.11 | 0.1 | 0.06 | 0.06 | 0.11 | 0.12 | 0.05 | 0.05 | 0.08 | 0.07 | 0.05 | 0.05 | 0.08 | 0.09 | 0.06
7 | 28 | E02 | 0.08 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.06 | 0.05 | 0.08 | 0.08 | 0.06 | 0.07 | 0.07 | 0.05 | 0.07
8 | 29 | E02 | 0.11 | 0.1 | 0.09 | 0.09 | 0.09 | 0.09 | 0.1 | 0.09 | 0.1 | 0.08 | 0.09 | 0.09 | 0.1 | 0.09 | 0.09
9 | 17 | E02 | 0.06 | 0.09 | 0.05 | 0.07 | 0.11 | 0.09 | 0.11 | 0.09 | 0.09 | 0.06 | 0.13 | 0.1 | 0.07 | 0.1 | 0.1
9 | 19 | F02 | 0.26 | 0.21 | 0.21 | 0.22 | 0.24 | 0.21 | 0.22 | 0.2 | 0.23 | 0.21 | 0.21 | 0.21 | 0.23 | 0.22 | 0.22
1 | 30 | E03 | 0.09 | 0.11 | 0.08 | 0.1 | 0.12 | 0.12 | 0.1 | 0.1 | 0.12 | 0.13 | 0.09 | 0.1 | 0.1 | 0.11 | 0.11
2 | 16 | E03 | 0.08 | 0.07 | 0.09 | 0.1 | 0.07 | 0.07 | 0.11 | 0.08 | 0.1 | 0.08 | 0.1 | 0.07 | 0.09 | 0.08 | 0.09
3 | 11 | E03 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.06 | 0.07 | 0.06 | 0.09 | 0.08 | 0.08 | 0.09 | 0.08 | 0.07 | 0.09
4 | 8 | E03 | 0.13 | 0.13 | 0.05 | 0.05 | 0.08 | 0.1 | 0.05 | 0.05 | 0.04 | 0.07 | 0.06 | 0.05 | 0.09 | 0.08 | 0.06
5 | 24 | E03 | 0.14 | 0.12 | 0.1 | 0.09 | 0.08 | 0.07 | 0.07 | 0.06 | 0.14 | 0.11 | 0.09 | 0.08 | 0.11 | 0.07 | 0.11
6 | 5 | E03 | 0.07 | 0.06 | 0.06 | 0.05 | 0.06 | 0.05 | 0.05 | 0.04 | 0.13 | 0.09 | 0.05 | 0.05 | 0.06 | 0.05 | 0.08
7 | 31 | E03 | 0.28 | 0.26 | 0.25 | 0.26 | 0.26 | 0.24 | 0.25 | 0.24 | 0.27 | 0.26 | 0.26 | 0.25 | 0.26 | 0.25 | 0.26 | Accidental tile swap. Data removed
8 | 31 | E03 | 0.1 | 0.1 | 0.09 | 0.08 | 0.1 | 0.08 | 0.09 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.09 | 0.09 | 0.08
9 | 2 | E03 | 0.12 | 0.07 | 0.06 | 0.05 | 0.08 | 0.05 | 0.07 | 0.05 | 0.08 | 0.07 | 0.1 | 0.06 | 0.08 | 0.07 | 0.08
1 | 7 | E04 | 0.07 | 0.09 | 0.11 | 0.07 | 0.08 | 0.09 | 0.08 | 0.07 | 0.1 | 0.1 | 0.06 | 0.08 | 0.09 | 0.08 | 0.09
2 | 29 | E04 | 0.08 | 0.06 | 0.08 | 0.07 | 0.07 | 0.06 | 0.12 | 0.06 | 0.07 | 0.06 | 0.1 | 0.06 | 0.07 | 0.08 | 0.07
3 | 35 | E04 | 0.08 | 0.07 | 0.07 | 0.06 | 0.09 | 0.07 | 0.07 | 0.06 | 0.09 | 0.06 | 0.07 | 0.06 | 0.07 | 0.08 | 0.07
4 | 20 | E04 | 0.09 | 0.08 | 0.07 | 0.06 | 0.11 | 0.09 | 0.05 | 0.06 | 0.11 | 0.12 | 0.06 | 0.06 | 0.08 | 0.08 | 0.09
5 | 23 | E04 | 0.12 | 0.1 | 0.07 | 0.07 | 0.11 | 0.09 | 0.09 | 0.07 | 0.13 | 0.11 | 0.1 | 0.07 | 0.09 | 0.1 | 0.1
6 | 31 | E04 | 0.09 | 0.08 | 0.07 | 0.06 | 0.12 | 0.07 | 0.07 | 0.06 | 0.11 | 0.08 | 0.07 | 0.06 | 0.08 | 0.09 | 0.08
7 | 11 | E04 | 0.06 | 0.05 | 0.05 | 0.05 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.07 | 0.06
8 | 16 | E04 | 0.1 | 0.1 | 0.09 | 0.08 | 0.09 | 0.07 | 0.1 | 0.08 | 0.09 | 0.08 | 0.07 | 0.08 | 0.09 | 0.09 | 0.08
9 | 1 | E04 | 0.17 | 0.05 | 0.07 | 0.04 | 0.08 | 0.07 | 0.05 | 0.05 | 0.16 | 0.07 | 0.06 | 0.05 | 0.08 | 0.07 | 0.09
1 | 26 | E05 | 0.09 | 0.11 | 0.08 | 0.09 | 0.08 | 0.1 | 0.08 | 0.08 | 0.1 | 0.11 | 0.08 | 0.1 | 0.09 | 0.09 | 0.1
2 | 32 | E05 | 0.1 | 0.06 | 0.08 | 0.06 | 0.07 | 0.06 | 0.09 | 0.06 | 0.11 | 0.06 | 0.09 | 0.06 | 0.08 | 0.07 | 0.08
3 | 23 | E05 | 0.1 | 0.09 | 0.07 | 0.06 | 0.09 | 0.08 | 0.07 | 0.07 | 0.1 | 0.08 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08
4 | 1 | E05 | 0.1 | 0.08 | 0.07 | 0.07 | 0.09 | 0.08 | 0.07 | 0.06 | 0.11 | 0.09 | 0.06 | 0.06 | 0.08 | 0.08 | 0.08
5 | 19 | E05 | 0.09 | 0.08 | 0.07 | 0.06 | 0.14 | 0.09 | 0.07 | 0.06 | 0.1 | 0.08 | 0.08 | 0.06 | 0.08 | 0.1 | 0.08
6 | 9 | E05 | 0.09 | 0.1 | 0.05 | 0.05 | 0.07 | 0.05 | 0.05 | 0.04 | 0.06 | 0.06 | 0.04 | 0.04 | 0.07 | 0.06 | 0.05
7 | 17 | E05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05
8 | 11 | E05 | 0.09 | 0.09 | 0.08 | 0.09 | 0.07 | 0.07 | 0.07 | 0.07 | 0.08 | 0.08 | 0.07 | 0.07 | 0.09 | 0.07 | 0.08
9 | 4 | E05 | 0.1 | 0.07 | 0.08 | 0.08 | 0.12 | 0.06 | 0.11 | 0.07 | 0.1 | 0.06 | 0.06 | 0.07 | 0.08 | 0.1 | 0.07
1 | 1 | E06 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.06 | 0.05 | 0.04
2 | 13 | E06 | 0.07 | 0.08 | 0.1 | 0.07 | 0.06 | 0.07 | 0.07 | 0.07 | 0.06 | 0.06 | 0.08 | 0.06 | 0.08 | 0.07 | 0.07
3 | 36 | E06 | 0.08 | 0.08 | 0.08 | 0.08 | 0.07 | 0.08 | 0.07 | 0.06 | 0.07 | 0.07 | 0.08 | 0.05 | 0.08 | 0.07 | 0.07
4 | 23 | E06 | 0.06 | 0.07 | 0.05 | 0.05 | 0.1 | 0.1 | 0.08 | 0.07 | 0.12 | 0.09 | 0.06 | 0.06 | 0.06 | 0.09 | 0.08
5 | 5 | E06 | 0.11 | 0.09 | 0.09 | 0.07 | 0.07 | 0.06 | 0.06 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.09 | 0.06 | 0.05
6 | 10 | E06 | 0.09 | 0.09 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07 | 0.06 | 0.08 | 0.06 | 0.06 | 0.05 | 0.08 | 0.07 | 0.06
7 | 20 | E06 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.06 | 0.06 | 0.07 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.06 | 0.06
8 | 20 | E06 | 0.06 | 0.06 | 0.06 | 0.05 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.11 | 0.08 | 0.09 | 0.06 | 0.07 | 0.09
9 | 13 | E06 | 0.08 | 0.09 | 0.11 | 0.08 | 0.13 | 0.1 | 0.12 | 0.06 | 0.14 | 0.08 | 0.11 | 0.07 | 0.09 | 0.12 | 0.1
1 | 10 | E07 | 0.09 | 0.11 | 0.07 | 0.07 | 0.08 | 0.11 | 0.07 | 0.08 | 0.15 | 0.14 | 0.07 | 0.08 | 0.09 | 0.09 | 0.11
2 | 2 | E07 | 0.07 | 0.06 | 0.11 | 0.08 | 0.06 | 0.06 | 0.08 | 0.08 | 0.05 | 0.06 | 0.13 | 0.06 | 0.08 | 0.07 | 0.08
3 | 3 | E07 | 0.07 | 0.05 | 0.06 | 0.05 | 0.06 | 0.04 | 0.06 | 0.05 | 0.06 | 0.05 | 0.07 | 0.05 | 0.06 | 0.05 | 0.06
4 | 18 | E07 | 0.13 | 0.12 | 0.06 | 0.05 | 0.08 | 0.08 | 0.07 | 0.06 | 0.1 | 0.06 | 0.06 | 0.05 | 0.09 | 0.08 | 0.07
5 | 11 | E07 | 0.11 | 0.08 | 0.08 | 0.08 | 0.12 | 0.08 | 0.08 | 0.06 | 0.11 | 0.09 | 0.08 | 0.06 | 0.09 | 0.09 | 0.09
6 | 22 | E07 | 0.07 | 0.07 | 0.06 | 0.06 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 0.05 | 0.07 | 0.06 | 0.06
7 | 16 | E07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.07 | 0.07 | 0.06 | 0.07 | 0.05 | 0.06 | 0.05 | 0.05 | 0.06 | 0.07 | 0.05
8 | 18 | E07 | 0.1 | 0.09 | 0.08 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.06 | 0.06 | 0.06 | 0.09 | 0.07 | 0.06
9 | 28 | E07 | 0.13 | 0.14 | 0.1 | 0.13 | 0.09 | 0.11 | 0.13 | 0.08 | 0.17 | 0.17 | 0.15 | 0.11 | 0.13 | 0.11 | 0.15
1 | 17 | E08 | 0.11 | 0.13 | 0.1 | 0.11 | 0.1 | 0.12 | 0.1 | 0.1 | 0.12 | 0.14 | 0.1 | 0.1 | 0.11 | 0.11 | 0.12
2 | 7 | E08 | 0.07 | 0.07 | 0.09 | 0.08 | 0.07 | 0.07 | 0.1 | 0.08 | 0.08 | 0.07 | 0.1 | 0.08 | 0.08 | 0.08 | 0.08
3 | 18 | E08 | 0.1 | 0.08 | 0.07 | 0.08 | 0.1 | 0.07 | 0.08 | 0.07 | 0.11 | 0.08 | 0.07 | 0.06 | 0.08 | 0.08 | 0.08
4 | 24 | E08 | 0.06 | 0.05 | 0.06 | 0.05 | 0.16 | 0.09 | 0.05 | 0.05 | 0.13 | 0.1 | 0.05 | 0.05 | 0.06 | 0.1 | 0.08
5 | 10 | E08 | 0.09 | 0.06 | 0.06 | 0.06 | 0.12 | 0.08 | 0.07 | 0.06 | 0.12 | 0.1 | 0.09 | 0.07 | 0.07 | 0.09 | 0.1
6 | 2 | E08 | 0.07 | 0.05 | 0.04 | 0.04 | 0.12 | 0.06 | 0.05 | 0.04 | 0.17 | 0.08 | 0.05 | 0.04 | 0.05 | 0.08 | 0.09
7 | 36 | E08 | 0.05 | 0.05 | 0.06 | 0.05 | 0.08 | 0.07 | 0.07 | 0.06 | 0.08 | 0.08 | 0.07 | 0.06 | 0.05 | 0.07 | 0.07
8 | 9 | E08 | 0.1 | 0.09 | 0.09 | 0.08 | 0.08 | 0.07 | 0.1 | 0.08 | 0.09 | 0.08 | 0.08 | 0.07 | 0.09 | 0.08 | 0.08
9 | 33 | E08 | 0.12 | 0.08 | 0.1 | 0.11 | 0.12 | 0.12 | 0.15 | 0.12 | 0.1 | 0.09 | 0.1 | 0.12 | 0.1 | 0.13 | 0.1
1 | 12 | E09 | 0.12 | 0.13 | 0.08 | 0.11 | 0.1 | 0.13 | 0.07 | 0.1 | 0.11 | 0.12 | 0.08 | 0.11 | 0.11 | 0.1 | 0.11
2 | 6 | E09 | 0.06 | 0.06 | 0.07 | 0.06 | 0.07 | 0.06 | 0.12 | 0.08 | 0.07 | 0.06 | 0.11 | 0.07 | 0.06 | 0.08 | 0.08
3 | 7 | E09 | 0.11 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.06 | 0.12 | 0.11 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.08
4 | 16 | E09 | 0.11 | 0.11 | 0.06 | 0.07 | 0.09 | 0.09 | 0.05 | 0.06 | 0.06 | 0.07 | 0.05 | 0.06 | 0.09 | 0.08 | 0.06
5 | 13 | E09 | 0.12 | 0.09 | 0.07 | 0.07 | 0.09 | 0.07 | 0.06 | 0.06 | 0.12 | 0.11 | 0.08 | 0.07 | 0.09 | 0.07 | 0.1
6 | 1 | E09 | 0.08 | 0.08 | 0.09 | 0.09 | 0.1 | 0.09 | 0.06 | 0.05 | 0.07 | 0.06 | 0.05 | 0.04 | 0.09 | 0.08 | 0.06
7 | 19 | E09 | 0.05 | 0.04 | 0.05 | 0.05 | 0.06 | 0.05 | 0.06 | 0.06 | 0.07 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06
8 | 22 | E09 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.06 | 0.07 | 0.06 | 0.09 | 0.12 | 0.08 | 0.07 | 0.08 | 0.07 | 0.09
9 | 3 | E09 | 0.14 | 0.04 | 0.07 | 0.06 | 0.13 | 0.06 | 0.08 | 0.07 | 0.12 | 0.08 | 0.07 | 0.05 | 0.08 | 0.09 | 0.08
1 | 27 | E10 | 0.11 | 0.1 | 0.09 | 0.1 | 0.08 | 0.09 | 0.06 | 0.09 | 0.1 | 0.11 | 0.07 | 0.12 | 0.1 | 0.08 | 0.1
2 | 26 | E10 | 0.07 | 0.08 | 0.11 | 0.07 | 0.07 | 0.07 | 0.14 | 0.09 | 0.08 | 0.07 | 0.11 | 0.07 | 0.08 | 0.09 | 0.08
3 | 2 | E10 | 0.09 | 0.07 | 0.07 | 0.07 | 0.05 | 0.05 | 0.05 | 0.04 | 0.08 | 0.05 | 0.05 | 0.04 | 0.08 | 0.05 | 0.06
4 | 12 | E10 | 0.06 | 0.06 | 0.05 | 0.06 | 0.07 | 0.07 | 0.04 | 0.05 | 0.1 | 0.1 | 0.06 | 0.06 | 0.06 | 0.06 | 0.08
5 | 7 | E10 | 0.12 | 0.09 | 0.06 | 0.06 | 0.14 | 0.09 | 0.07 | 0.05 | 0.1 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1 | 0.08
6 | 20 | E10 | 0.1 | 0.08 | 0.06 | 0.06 | 0.09 | 0.07 | 0.06 | 0.05 | 0.1 | 0.08 | 0.06 | 0.05 | 0.08 | 0.07 | 0.07
7 | 33 | E10 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06 | 0.06 | 0.05 | 0.05 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06 | 0.06
8 | 35 | E10 | 0.09 | 0.08 | 0.09 | 0.08 | 0.06 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.09 | 0.07 | 0.07
9 | 21 | E10 | 0.11 | 0.07 | 0.12 | 0.08 | 0.08 | 0.08 | 0.12 | 0.16 | 0.1 | 0.05 | 0.12 | 0.07 | 0.1 | 0.09 | 0.09
1 | 18 | E11 | 0.09 | 0.11 | 0.08 | 0.09 | 0.11 | 0.13 | 0.1 | 0.12 | 0.1 | 0.12 | 0.1 | 0.13 | 0.09 | 0.11 | 0.11
2 | 17 | E11 | 0.06 | 0.08 | 0.12 | 0.06 | 0.05 | 0.06 | 0.09 | 0.06 | 0.06 | 0.05 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07
3 | 30 | E11 | 0.08 | 0.07 | 0.06 | 0.06 | 0.1 | 0.08 | 0.07 | 0.07 | 0.07 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | 0.07
4 | 21 | E11 | 0.08 | 0.09 | 0.06 | 0.05 | 0.13 | 0.1 | 0.05 | 0.05 | 0.11 | 0.1 | 0.06 | 0.06 | 0.07 | 0.09 | 0.08
5 | 9 | E11 | 0.11 | 0.08 | 0.07 | 0.07 | 0.11 | 0.08 | 0.06 | 0.05 | 0.11 | 0.11 | 0.07 | 0.06 | 0.08 | 0.08 | 0.09
6 | 27 | E11 | 0.1 | 0.1 | 0.06 | 0.07 | 0.07 | 0.06 | 0.05 | 0.06 | 0.08 | 0.08 | 0.05 | 0.05 | 0.08 | 0.06 | 0.07
7 | 5 | E11 | 0.07 | 0.06 | 0.1 | 0.06 | 0.1 | 0.07 | 0.06 | 0.06 | 0.09 | 0.06 | 0.06 | 0.05 | 0.07 | 0.08 | 0.07
8 | 8 | E11 | 0.1 | 0.11 | 0.08 | 0.09 | 0.06 | 0.06 | 0.06 | 0.06 | 0.07 | 0.06 | 0.06 | 0.06 | 0.1 | 0.06 | 0.06
9 | 11 | E11 | 0.05 | 0.05 | 0.13 | 0.12 | 0.1 | 0.07 | 0.14 | 0.09 | 0.1 | 0.06 | 0.12 | 0.07 | 0.09 | 0.1 | 0.09
1 | 14 | E12 | 0.08 | 0.11 | 0.09 | 0.09 | 0.1 | 0.12 | 0.09 | 0.1 | 0.09 | 0.12 | 0.09 | 0.1 | 0.09 | 0.1 | 0.1
2 | 23 | E12 | 0.07 | 0.09 | 0.08 | 0.06 | 0.11 | 0.06 | 0.1 | 0.07 | 0.06 | 0.06 | 0.1 | 0.07 | 0.08 | 0.09 | 0.07
3 | 8 | E12 | 0.07 | 0.06 | 0.06 | 0.05 | 0.05 | 0.05 | 0.04 | 0.05 | 0.1 | 0.06 | 0.07 | 0.05 | 0.06 | 0.05 | 0.07
4 | 11 | E12 | 0.15 | 0.11 | 0.05 | 0.05 | 0.1 | 0.08 | 0.05 | 0.05 | 0.11 | 0.1 | 0.05 | 0.05 | 0.09 | 0.08 | 0.08
5 | 18 | E12 | 0.16 | 0.13 | 0.13 | 0.12 | 0.13 | 0.1 | 0.08 | 0.07 | 0.14 | 0.14 | 0.1 | 0.08 | 0.14 | 0.1 | 0.12
6 | 12 | E12 | 0.08 | 0.07 | 0.05 | 0.04 | 0.07 | 0.05 | 0.05 | 0.04 | 0.08 | 0.05 | 0.05 | 0.05 | 0.06 | 0.06 | 0.06
7 | 34 | E12 | 0.07 | 0.06 | 0.06 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 0.05 | 0.06 | 0.06 | 0.05 | 0.07 | 0.06 | 0.06
8 | 30 | E12 | 0.11 | 0.1 | 0.1 | 0.09 | 0.09 | 0.09 | 0.11 | 0.09 | 0.08 | 0.07 | 0.07 | 0.08 | 0.1 | 0.1 | 0.08
9 | 23 | E12 | 0.12 | 0.1 | 0.06 | 0.08 | 0.1 | 0.14 | 0.08 | 0.1 | 0.06 | 0.13 | 0.06 | 0.09 | 0.09 | 0.11 | 0.09
1 | 34 | F01 | 0.21 | 0.26 | 0.21 | 0.2 | 0.21 | 0.25 | 0.21 | 0.19 | 0.21 | 0.25 | 0.21 | 0.2 | 0.22 | 0.22 | 0.22
2 | 22 | F01 | 0.24 | 0.2 | 0.24 | 0.22 | 0.24 | 0.21 | 0.28 | 0.23 | 0.24 | 0.21 | 0.23 | 0.22 | 0.23 | 0.24 | 0.23
3 | 16 | F01 | 0.21 | 0.19 | 0.19 | 0.19 | 0.21 | 0.19 | 0.19 | 0.19 | 0.22 | 0.19 | 0.19 | 0.19 | 0.2 | 0.2 | 0.2
4 | 4 | F01 | 0.21 | 0.17 | 0.17 | 0.18 | 0.23 | 0.18 | 0.18 | 0.18 | 0.24 | 0.19 | 0.2 | 0.19 | 0.18 | 0.2 | 0.21
5 | 6 | F01 | 0.23 | 0.19 | 0.19 | 0.19 | 0.19 | 0.19 | 0.19 | 0.18 | 0.23 | 0.2 | 0.2 | 0.19 | 0.2 | 0.19 | 0.21
6 | 30 | F01 | 0.24 | 0.2 | 0.2 | 0.19 | 0.22 | 0.18 | 0.2 | 0.19 | 0.21 | 0.17 | 0.18 | 0.17 | 0.21 | 0.2 | 0.18
7 | 12 | F01 | 0.26 | 0.25 | 0.23 | 0.22 | 0.25 | 0.24 | 0.22 | 0.24 | 0.26 | 0.24 | 0.24 | 0.23 | 0.24 | 0.24 | 0.24
8 | 15 | F01 | 0.22 | 0.2 | 0.2 | 0.2 | 0.23 | 0.21 | 0.2 | 0.21 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21 | 0.21 | 0.2
9 | 27 | F01 | 0.23 | 0.18 | 0.2 | 0.2 | 0.23 | 0.18 | 0.19 | 0.2 | 0.23 | 0.19 | 0.21 | 0.2 | 0.2 | 0.2 | 0.21
1 | 35 | F02 | 0.22 | 0.27 | 0.2 | 0.21 | 0.22 | 0.26 | 0.21 | 0.21 | 0.23 | 0.26 | 0.21 | 0.21 | 0.23 | 0.23 | 0.23
2 | 10 | F02 | 0.19 | 0.19 | 0.22 | 0.2 | 0.2 | 0.21 | 0.23 | 0.21 | 0.21 | 0.22 | 0.23 | 0.22 | 0.2 | 0.21 | 0.22
3 | 1 | F02 | 0.19 | 0.19 | 0.18 | 0.19 | 0.25 | 0.24 | 0.25 | 0.24 | 0.33 | 0.2 | 0.21 | 0.21 | 0.19 | 0.25 | 0.24
4 | 17 | F02 | 0.25 | 0.24 | 0.22 | 0.23 | 0.25 | 0.23 | 0.22 | 0.22 | 0.27 | 0.23 | 0.22 | 0.22 | 0.24 | 0.23 | 0.24
5 | 16 | F02 | 0.24 | 0.22 | 0.21 | 0.2 | 0.22 | 0.19 | 0.2 | 0.21 | 0.22 | 0.2 | 0.2 | 0.21 | 0.22 | 0.2 | 0.21
6 | 7 | F02 | 0.22 | 0.22 | 0.22 | 0.21 | 0.22 | 0.2 | 0.19 | 0.19 | 0.22 | 0.2 | 0.19 | 0.18 | 0.22 | 0.2 | 0.2
7 | 29 | F02 | 0.24 | 0.23 | 0.24 | 0.24 | 0.25 | 0.22 | 0.21 | 0.21 | 0.21 | 0.21 | 0.22 | 0.23 | 0.24 | 0.23 | 0.22
8 | 27 | F02 | 0.21 | 0.22 | 0.22 | 0.21 | 0.23 | 0.21 | 0.22 | 0.21 | 0.21 | 0.21 | 0.22 | 0.21 | 0.22 | 0.22 | 0.21
1 | 4 | F03 | 0.2 | 0.22 | 0.19 | 0.18 | 0.2 | 0.21 | 0.2 | 0.19 | 0.21 | 0.2 | 0.2 | 0.19 | 0.2 | 0.2 | 0.2
2 | 36 | F03 | 0.25 | 0.2 | 0.22 | 0.21 | 0.23 | 0.21 | 0.27 | 0.22 | 0.31 | 0.21 | 0.23 | 0.22 | 0.22 | 0.24 | 0.24
3 | 20 | F03 | 0.2 | 0.19 | 0.19 | 0.19 | 0.19 | 0.17 | 0.19 | 0.19 | 0.19 | 0.17 | 0.18 | 0.18 | 0.19 | 0.18 | 0.18
4 | 14 | F03 | 0.24 | 0.22 | 0.21 | 0.2 | 0.26 | 0.21 | 0.2 | 0.2 | 0.25 | 0.2 | 0.2 | 0.21 | 0.22 | 0.22 | 0.22
5 | 15 | F03 | 0.22 | 0.18 | 0.19 | 0.18 | 0.21 | 0.18 | 0.19 | 0.18 | 0.22 | 0.18 | 0.19 | 0.18 | 0.19 | 0.19 | 0.19
6 | 28 | F03 | 0.22 | 0.18 | 0.18 | 0.18 | 0.22 | 0.18 | 0.18 | 0.18 | 0.22 | 0.18 | 0.17 | 0.17 | 0.19 | 0.19 | 0.19
7 | 21 | F03 | 0.23 | 0.21 | 0.21 | 0.2 | 0.21 | 0.19 | 0.21 | 0.2 | 0.22 | 0.2 | 0.2 | 0.2 | 0.21 | 0.2 | 0.21
8 | 28 | F03 | 0.21 | 0.19 | 0.19 | 0.2 | 0.21 | 0.19 | 0.19 | 0.2 | 0.22 | 0.2 | 0.19 | 0.2 | 0.2 | 0.2 | 0.2
9 | 8 | F03 | 0.25 | 0.19 | 0.2 | 0.2 | 0.25 | 0.19 | 0.2 | 0.2 | 0.24 | 0.19 | 0.2 | 0.2 | 0.21 | 0.21 | 0.21
1 | 13 | F04 | 0.23 | 0.26 | 0.23 | 0.22 | 0.23 | 0.24 | 0.23 | 0.22 | 0.22 | 0.26 | 0.22 | 0.21 | 0.24 | 0.23 | 0.23
2 | 4 | F04 | 0.17 | 0.17 | 0.19 | 0.19 | 0.19 | 0.18 | 0.21 | 0.2 | 0.19 | 0.19 | 0.23 | 0.21 | 0.18 | 0.19 | 0.21
3 | 33 | F04 | 0.2 | 0.2 | 0.18 | 0.19 | 0.19 | 0.18 | 0.19 | 0.19 | 0.2 | 0.18 | 0.18 | 0.19 | 0.19 | 0.19 | 0.19
4 | 5 | F04 | 0.25 | 0.22 | 0.21 | 0.24 | 0.27 | 0.22 | 0.21 | 0.23 | 0.27 | 0.22 | 0.21 | 0.24 | 0.23 | 0.23 | 0.24
5 | 30 | F04 | 0.22 | 0.21 | 0.2 | 0.21 | 0.2 | 0.22 | 0.2 | 0.21 | 0.29 | 0.2 | 0.21 | 0.21 | 0.21 | 0.21 | 0.23
6 | 24 | F04 | 0.24 | 0.2 | 0.19 | 0.2 | 0.22 | 0.19 | 0.18 | 0.19 | 0.22 | 0.19 | 0.18 | 0.19 | 0.21 | 0.2 | 0.2
7 | 22 | F04 | 0.26 | 0.24 | 0.24 | 0.25 | 0.24 | 0.22 | 0.21 | 0.22 | 0.25 | 0.24 | 0.23 | 0.22 | 0.25 | 0.22 | 0.24
8 | 26 | F04 | 0.2 | 0.2 | 0.2 | 0.2 | 0.21 | 0.21 | 0.21 | 0.21 | 0.22 | 0.22 | 0.22 | 0.22 | 0.2 | 0.21 | 0.22
9 | 15 | F04 | 0.25 | 0.21 | 0.21 | 0.22 | 0.23 | 0.21 | 0.21 | 0.22 | 0.23 | 0.21 | 0.21 | 0.22 | 0.22 | 0.22 | 0.22
1 | 2 | F05 | 0.26 | 0.25 | 0.22 | 0.24 | 0.25 | 0.26 | 0.24 | 0.25 | 0.26 | 0.25 | 0.24 | 0.26 | 0.24 | 0.25 | 0.25
2 | 1 | F05 | 0.19 | 0.2 | 0.22 | 0.2 | 0.21 | 0.22 | 0.24 | 0.22 | 0.22 | 0.22 | 0.25 | 0.22 | 0.2 | 0.22 | 0.23
3 | 5 | F05 | 0.2 | 0.24 | 0.24 | 0.25 | 0.24 | 0.24 | 0.22 | 0.23 | 0.24 | 0.24 | 0.22 | 0.22 | 0.23 | 0.23 | 0.23
4 | 22 | F05 | 0.27 | 0.26 | 0.28 | 0.26 | 0.27 | 0.26 | 0.27 | 0.25 | 0.27 | 0.26 | 0.27 | 0.26 | 0.27 | 0.27 | 0.27
5 | 29 | F05 | 0.26 | 0.24 | 0.25 | 0.24 | 0.24 | 0.23 | 0.23 | 0.24 | 0.24 | 0.22 | 0.23 | 0.23 | 0.25 | 0.23 | 0.23
6 | 23 | F05 | 0.25 | 0.23 | 0.22 | 0.22 | 0.24 | 0.23 | 0.22 | 0.21 | 0.25 | 0.23 | 0.21 | 0.21 | 0.23 | 0.23 | 0.23
7 | 27 | F05 | 0.27 | 0.26 | 0.24 | 0.23 | 0.26 | 0.26 | 0.25 | 0.25 | 0.28 | 0.24 | 0.26 | 0.26 | 0.25 | 0.26 | 0.26
8 | 4 | F05 | 0.25 | 0.27 | 0.27 | 0.25 | 0.25 | 0.26 | 0.26 | 0.26 | 0.25 | 0.26 | 0.26 | 0.25 | 0.26 | 0.26 | 0.26
9 | 25 | F05 | 0.25 | 0.26 | 0.26 | 0.25 | 0.24 | 0.25 | 0.24 | 0.25 | 0.26 | 0.23 | 0.25 | 0.25 | 0.26 | 0.24 | 0.25
1 | 5 | F06 | 0.25 | 0.23 | 0.26 | 0.24 | 0.26 | 0.24 | 0.26 | 0.24 | 0.26 | 0.24 | 0.27 | 0.24 | 0.25 | 0.25 | 0.25
2 | 18 | F06 | 0.25 | 0.24 | 0.26 | 0.26 | 0.26 | 0.24 | 0.25 | 0.25 | 0.27 | 0.25 | 0.26 | 0.27 | 0.25 | 0.25 | 0.26
3 | 27 | F06 | 0.24 | 0.23 | 0.22 | 0.26 | 0.24 | 0.22 | 0.23 | 0.24 | 0.25 | 0.22 | 0.22 | 0.23 | 0.24 | 0.23 | 0.23
4 | 31 | F06 | 0.29 | 0.24 | 0.25 | 0.28 | 0.28 | 0.24 | 0.26 | 0.28 | 0.28 | 0.25 | 0.25 | 0.26 | 0.27 | 0.26 | 0.26
5 | 35 | F06 | 0.25 | 0.21 | 0.22 | 0.25 | 0.25 | 0.22 | 0.23 | 0.24 | 0.24 | 0.22 | 0.22 | 0.24 | 0.23 | 0.23 | 0.23
6 | 15 | F06 | 0.27 | 0.22 | 0.22 | 0.23 | 0.25 | 0.22 | 0.22 | 0.23 | 0.26 | 0.22 | 0.22 | 0.24 | 0.24 | 0.23 | 0.24
7 | 1 | F06 | 0.25 | 0.23 | 0.22 | 0.23 | 0.28 | 0.24 | 0.24 | 0.25 | 0.28 | 0.25 | 0.24 | 0.25 | 0.23 | 0.25 | 0.26
8 | 10 | F06 | 0.26 | 0.24 | 0.24 | 0.25 | 0.25 | 0.24 | 0.25 | 0.26 | 0.26 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25 | 0.25
9 | 22 | F06 | 0.27 | 0.23 | 0.24 | 0.24 | 0.26 | 0.23 | 0.24 | 0.23 | 0.25 | 0.23 | 0.24 | 0.24 | 0.25 | 0.24 | 0.24
1 | 11 | F07 | 0.26 | 0.27 | 0.25 | 0.23 | 0.25 | 0.26 | 0.25 | 0.24 | 0.24 | 0.27 | 0.25 | 0.24 | 0.25 | 0.25 | 0.25
2 | 35 | F07 | 0.27 | 0.26 | 0.27 | 0.27 | 0.28 | 0.24 | 0.26 | 0.26 | 0.28 | 0.25 | 0.26 | 0.26 | 0.27 | 0.26 | 0.26
3 | 14 | F07 | 0.24 | 0.22 | 0.22 | 0.24 | 0.24 | 0.23 | 0.21 | 0.23 | 0.29 | 0.23 | 0.22 | 0.23 | 0.23 | 0.23 | 0.24
4 | 27 | F07 | 0.3 | 0.27 | 0.26 | 0.27 | 0.3 | 0.26 | 0.25 | 0.27 | 0.3 | 0.26 | 0.26 | 0.27 | 0.28 | 0.27 | 0.27
5 | 2 | F07 | 0.21 | 0.21 | 0.2 | 0.22 | 0.23 | 0.22 | 0.2 | 0.23 | 0.22 | 0.23 | 0.21 | 0.24 | 0.21 | 0.22 | 0.23
6 | 4 | F07 | 0.28 | 0.25 | 0.24 | 0.24 | 0.25 | 0.23 | 0.22 | 0.23 | 0.24 | 0.22 | 0.21 | 0.22 | 0.25 | 0.23 | 0.22
7 | 35 | F07 | 0.29 | 0.27 | 0.26 | 0.27 | 0.28 | 0.26 | 0.26 | 0.26 | 0.27 | 0.27 | 0.25 | 0.27 | 0.27 | 0.27 | 0.27
8 | 14 | F07 | 0.25 | 0.24 | 0.25 | 0.25 | 0.25 | 0.24 | 0.26 | 0.26 | 0.26 | 0.25 | 0.24 | 0.24 | 0.25 | 0.25 | 0.25
9 | 6 | F07 | 0.35 | 0.24 | 0.24 | 0.24 | 0.28 | 0.24 | 0.24 | 0.25 | 0.26 | 0.24 | 0.24 | 0.24 | 0.27 | 0.25 | 0.25
1 | 21 | F08 | 0.24 | 0.27 | 0.23 | 0.22 | 0.23 | 0.26 | 0.24 | 0.22 | 0.23 | 0.26 | 0.24 | 0.24 | 0.24 | 0.24 | 0.24
2 | 21 | F08 | 0.23 | 0.21 | 0.25 | 0.21 | 0.23 | 0.21 | 0.25 | 0.22 | 0.23 | 0.21 | 0.25 | 0.21 | 0.23 | 0.23 | 0.23
3 | 22 | F08 | 0.22 | 0.21 | 0.23 | 0.2 | 0.23 | 0.21 | 0.22 | 0.2 | 0.23 | 0.2 | 0.22 | 0.2 | 0.22 | 0.22 | 0.21
4 | 13 | F08 | 0.28 | 0.24 | 0.25 | 0.24 | 0.27 | 0.24 | 0.23 | 0.21 | 0.24 | 0.22 | 0.23 | 0.21 | 0.25 | 0.25 | 0.23
5 | 3 | F08 | 0.22 | 0.19 | 0.19 | 0.19 | 0.22 | 0.2 | 0.21 | 0.2 | 0.22 | 0.2 | 0.21 | 0.2 | 0.2 | 0.21 | 0.21
6 | 11 | F08 | 0.24 | 0.21 | 0.23 | 0.2 | 0.24 | 0.2 | 0.2 | 0.19 | 0.24 | 0.2 | 0.21 | 0.19 | 0.22 | 0.21 | 0.21
7 | 8 | F08 | 0.26 | 0.23 | 0.24 | 0.23 | 0.26 | 0.23 | 0.23 | 0.23 | 0.26 | 0.23 | 0.24 | 0.23 | 0.24 | 0.24 | 0.24
8 | 25 | F08 | 0.24 | 0.21 | 0.22 | 0.2 | 0.23 | 0.2 | 0.23 | 0.21 | 0.23 | 0.21 | 0.23 | 0.2 | 0.22 | 0.22 | 0.22
9 | 26 | F08 | 0.24 | 0.21 | 0.24 | 0.21 | 0.25 | 0.21 | 0.22 | 0.24 | 0.23 | 0.22 | 0.23 | 0.2 | 0.23 | 0.23 | 0.22
1 | 20 | F09 | 0.28 | 0.28 | 0.28 | 0.27 | 0.26 | 0.29 | 0.26 | 0.24 | 0.27 | 0.29 | 0.27 | 0.26 | 0.28 | 0.27 | 0.27
2 | 5 | F09 | 0.23 | 0.22 | 0.24 | 0.24 | 0.24 | 0.23 | 0.24 | 0.24 | 0.24 | 0.23 | 0.25 | 0.24 | 0.23 | 0.24 | 0.24
3 | 6 | F09 | 0.26 | 0.25 | 0.25 | 0.26 | 0.27 | 0.25 | 0.24 | 0.25 | 0.26 | 0.25 | 0.25 | 0.25 | 0.26 | 0.25 | 0.25
4 | 3 | F09 | 0.27 | 0.23 | 0.22 | 0.23 | 0.25 | 0.24 | 0.24 | 0.24 | 0.28 | 0.24 | 0.23 | 0.22 | 0.24 | 0.24 | 0.24
5 | 22 | F09 | 0.27 | 0.05 | 0.25 | 0.26 | 0.27 | 0.24 | 0.24 | 0.25 | 0.26 | 0.23 | 0.24 | 0.25 | 0.21 | 0.25 | 0.25
6 | 21 | F09 | 0.28 | 0.25 | 0.25 | 0.25 | 0.25 | 0.22 | 0.23 | 0.24 | 0.26 | 0.22 | 0.23 | 0.24 | 0.26 | 0.23 | 0.24
7 | 24 | F09 | 0.29 | 0.27 | 0.25 | 0.26 | 0.28 | 0.25 | 0.27 | 0.27 | 0.28 | 0.27 | 0.25 | 0.25 | 0.27 | 0.27 | 0.26
8 | 17 | F09 | 0.27 | 0.25 | 0.24 | 0.25 | 0.26 | 0.24 | 0.26 | 0.26 | 0.29 | 0.26 | 0.26 | 0.25 | 0.25 | 0.25 | 0.27
9 | 31 | F09 | 0.28 | 0.24 | 0.25 | 0.24 | 0.29 | 0.24 | 0.24 | 0.24 | 0.27 | 0.23 | 0.23 | 0.24 | 0.25 | 0.26 | 0.24
1 | 31 | F10 | 0.24 | 0.27 | 0.22 | 0.22 | 0.24 | 0.25 | 0.22 | 0.23 | 0.23 | 0.25 | 0.22 | 0.23 | 0.24 | 0.24 | 0.23
2 | 14 | F10 | 0.22 | 0.22 | 0.24 | 0.23 | 0.22 | 0.22 | 0.23 | 0.23 | 0.23 | 0.23 | 0.22 | 0.23 | 0.23 | 0.22 | 0.23
3 | 26 | F10 | 0.24 | 0.22 | 0.21 | 0.22 | 0.22 | 0.22 | 0.2 | 0.22 | 0.22 | 0.21 | 0.21 | 0.21 | 0.22 | 0.21 | 0.21
4 | 30 | F10 | 0.24 | 0.24 | 0.24 | 0.25 | 0.25 | 0.22 | 0.22 | 0.24 | 0.24 | 0.23 | 0.22 | 0.23 | 0.24 | 0.23 | 0.23
5 | 36 | F10 | 0.22 | 0.18 | 0.19 | 0.17 | 0.21 | 0.19 | 0.19 | 0.2 | 0.22 | 0.19 | 0.2 | 0.2 | 0.19 | 0.2 | 0.2
6 | 36 | F10 | 0.22 | 0.19 | 0.19 | 0.19 | 0.23 | 0.19 | 0.18 | 0.19 | 0.23 | 0.19 | 0.17 | 0.18 | 0.2 | 0.2 | 0.19
7 | 7 | F10 | 0.23 | 0.21 | 0.21 | 0.21 | 0.24 | 0.22 | 0.21 | 0.21 | 0.25 | 0.23 | 0.22 | 0.22 | 0.22 | 0.22 | 0.23
8 | 33 | F10 | 0.2 | 0.19 | 0.2 | 0.2 | 0.21 | 0.2 | 0.2 | 0.21 | 0.21 | 0.2 | 0.21 | 0.2 | 0.2 | 0.2 | 0.21
9 | 14 | F10 | 0.25 | 0.22 | 0.2 | 0.22 | 0.25 | 0.21 | 0.22 | 0.21 | 0.25 | 0.21 | 0.22 | 0.22 | 0.22 | 0.23 | 0.23
1 | 28 | F11 | 0.22 | 0.26 | 0.25 | 0.24 | 0.22 | 0.25 | 0.23 | 0.23 | 0.23 | 0.26 | 0.23 | 0.24 | 0.24 | 0.23 | 0.24
2 | 20 | F11 | 0.24 | 0.22 | 0.25 | 0.22 | 0.24 | 0.23 | 0.27 | 0.23 | 0.24 | 0.23 | 0.27 | 0.23 | 0.23 | 0.25 | 0.24
3 | 9 | F11 | 0.19 | 0.18 | 0.21 | 0.18 | 0.2 | 0.19 | 0.21 | 0.19 | 0.21 | 0.19 | 0.21 | 0.19 | 0.19 | 0.2 | 0.2
4 | 19 | F11 | 0.29 | 0.25 | 0.27 | 0.24 | 0.27 | 0.23 | 0.25 | 0.23 | 0.26 | 0.23 | 0.24 | 0.22 | 0.26 | 0.25 | 0.24
5 | 12 | F11 | 0.22 | 0.19 | 0.2 | 0.19 | 0.23 | 0.21 | 0.21 | 0.18 | 0.23 | 0.2 | 0.2 | 0.19 | 0.2 | 0.22 | 0.21
6 | 19 | F11 | 0.27 | 0.23 | 0.22 | 0.2 | 0.24 | 0.21 | 0.22 | 0.19 | 0.25 | 0.21 | 0.21 | 0.2 | 0.23 | 0.22 | 0.22
7 | 25 | F11 | 0.29 | 0.26 | 0.25 | 0.24 | 0.27 | 0.24 | 0.25 | 0.24 | 0.27 | 0.24 | 0.23 | 0.23 | 0.26 | 0.25 | 0.24
8 | 32 | F11 | 0.22 | 0.22 | 0.23 | 0.2 | 0.22 | 0.22 | 0.22 | 0.21 | 0.21 | 0.22 | 0.23 | 0.2 | 0.22 | 0.22 | 0.22
9 | 34 | F11 | 0.24 | 0.21 | 0.21 | 0.18 | 0.22 | 0.2 | 0.24 | 0.19 | 0.23 | 0.2 | 0.22 | 0.19 | 0.21 | 0.22 | 0.21
1 | 22 | F12 | 0.22 | 0.26 | 0.21 | 0.21 | 0.22 | 0.25 | 0.21 | 0.21 | 0.21 | 0.24 | 0.21 | 0.21 | 0.23 | 0.23 | 0.22
2 | 31 | F12 | 0.21 | 0.23 | 0.31 | 0.21 | 0.24 | 0.23 | 0.27 | 0.22 | 0.24 | 0.23 | 0.33 | 0.22 | 0.24 | 0.25 | 0.26
3 | 15 | F12 | 0.2 | 0.19 | 0.19 | 0.19 | 0.19 | 0.19 | 0.17 | 0.18 | 0.19 | 0.19 | 0.18 | 0.18 | 0.19 | 0.18 | 0.19
4 | 10 | F12 | 0.23 | 0.22 | 0.19 | 0.2 | 0.23 | 0.21 | 0.2 | 0.2 | 0.2 | 0.21 | 0.19 | 0.2 | 0.21 | 0.21 | 0.2
5 | 25 | F12 | 0.23 | 0.19 | 0.18 | 0.19 | 0.2 | 0.22 | 0.19 | 0.2 | 0.22 | 0.19 | 0.19 | 0.19 | 0.2 | 0.2 | 0.2
6 | 32 | F12 | 0.21 | 0.14 | 0.18 | 0.18 | 0.21 | 0.21 | 0.19 | 0.19 | 0.23 | 0.2 | 0.18 | 0.18 | 0.18 | 0.2 | 0.2
7 | 10 | F12 | 0.2 | 0.21 | 0.19 | 0.19 | 0.22 | 0.21 | 0.2 | 0.2 | 0.23 | 0.22 | 0.2 | 0.21 | 0.2 | 0.21 | 0.22
8 | 34 | F12 | 0.22 | 0.22 | 0.21 | 0.2 | 0.2 | 0.21 | 0.2 | 0.22 | 0.21 | 0.22 | 0.21 | 0.2 | 0.21 | 0.2 | 0.21
9 | 20 | F12 | 0.21 | 0.2 | 0.17 | 0.19 | 0.21 | 0.2 | 0.19 | 0.2 | 0.21 | 0.2 | 0.19 | 0.2 | 0.19 | 0.2 | 0.2
1 | 19 | G01 | 0.69 | 0.69 | 0.67 | 0.65 | 0.69 | 0.7 | 0.66 | 0.65 | 0.7 | 0.7 | 0.67 | 0.66 | 0.68 | 0.68 | 0.68
2 | 3 | G01 | 0.58 | 0.57 | 0.58 | 0.58 | 0.6 | 0.59 | 0.59 | 0.06 | 0.62 | 0.59 | 0.61 | 0.59 | 0.58 | 0.59 | 0.6
3 | 32 | G01 | 0.51 | 0.47 | 0.48 | 0.49 | 0.52 | 0.49 | 0.49 | 0.48 | 0.52 | 0.47 | 0.48 | 0.47 | 0.49 | 0.5 | 0.49
4 | 36 | G01 | 0.67 | 0.6 | 0.63 | 0.63 | 0.66 | 0.62 | 0.63 | 0.65 | 0.7 | 0.62 | 0.65 | 0.64 | 0.63 | 0.64 | 0.65
5 | 32 | G01 | 0.61 | 0.59 | 0.59 | 0.6 | 0.62 | 0.59 | 0.59 | 0.6 | 0.61 | 0.59 | 0.6 | 0.59 | 0.6 | 0.6 | 0.6
6 | 25 | G01 | 0.57 | 0.5 | 0.52 | 0.51 | 0.55 | 0.48 | 0.48 | 0.49 | 0.56 | 0.47 | 0.48 | 0.47 | 0.53 | 0.5 | 0.5
7 | 14 | G01 | 0.67 | 0.62 | 0.65 | 0.66 | 0.69 | 0.66 | 0.64 | 0.67 | 0.67 | 0.68 | 0.66 | 0.67 | 0.65 | 0.66 | 0.67
8 | 23 | G01 | 0.53 | 0.45 | 0.46 | 0.46 | 0.5 | 0.47 | 0.51 | 0.47 | 0.52 | 0.46 | 0.5 | 0.47 | 0.48 | 0.49 | 0.49
9 | 29 | G01 | 0.62 | 0.58 | 0.59 | 0.58 | 0.63 | 0.6 | 0.57 | 0.58 | 0.6 | 0.6 | 0.57 | 0.57 | 0.59 | 0.6 | 0.59
1 | 32 | G02 | 0.68 | 0.67 | 0.67 | 0.66 | 0.68 | 0.67 | 0.66 | 0.66 | 0.68 | 0.67 | 0.67 | 0.66 | 0.67 | 0.67 | 0.67
2 | 24 | G02 | 0.65 | 0.63 | 0.63 | 0.63 | 0.66 | 0.66 | 0.65 | 0.66 | 0.66 | 0.67 | 0.66 | 0.67 | 0.64 | 0.66 | 0.67
3 | 12 | G02 | 0.54 | 0.51 | 0.49 | 0.5 | 0.55 | 0.51 | 0.51 | 0.51 | 0.57 | 0.52 | 0.5 | 0.52 | 0.51 | 0.52 | 0.53
4 | 35 | G02 | 0.69 | 0.68 | 0.69 | 0.66 | 0.67 | 0.67 | 0.65 | 0.65 | 0.67 | 0.65 | 0.64 | 0.64 | 0.68 | 0.66 | 0.65
5 | 31 | G02 | 0.61 | 0.59 | 0.56 | 0.59 | 0.63 | 0.61 | 0.57 | 0.59 | 0.62 | 0.61 | 0.59 | 0.6 | 0.59 | 0.6 | 0.61
6 | 29 | G02 | 0.62 | 0.57 | 0.57 | 0.53 | 0.58 | 0.52 | 0.52 | 0.53 | 0.57 | 0.52 | 0.56 | 0.49 | 0.57 | 0.54 | 0.54
7 | 2 | G02 | 0.69 | 0.65 | 0.63 | 0.66 | 0.69 | 0.66 | 0.64 | 0.67 | 0.69 | 0.67 | 0.67 | 0.7 | 0.66 | 0.66 | 0.68
8 | 12 | G02 | 0.57 | 0.53 | 0.51 | 0.54 | 0.53 | 0.52 | 0.53 | 0.54 | 0.55 | 0.52 | 0.52 | 0.52 | 0.54 | 0.53 | 0.53
9 | 7 | G02 | 0.65 | 0.61 | 0.63 | 0.6 | 0.64 | 0.63 | 0.64 | 0.63 | 0.64 | 0.65 | 0.63 | 0.62 | 0.62 | 0.64 | 0.64
1 | 33 | G03 | 0.69 | 0.68 | 0.67 | 0.65 | 0.67 | 0.67 | 0.69 | 0.63 | 0.67 | 0.67 | 0.67 | 0.65 | 0.67 | 0.68 | 0.67
2 | 12 | G03 | 0.66 | 0.61 | 0.64 | 0.65 | 0.66 | 0.61 | 0.67 | 0.65 | 0.68 | 0.61 | 0.7 | 0.65 | 0.64 | 0.65 | 0.66
3 | 21 | G03 | 0.51 | 0.45 | 0.46 | 0.46 | 0.52 | 0.45 | 0.47 | 0.47 | 0.52 | 0.47 | 0.48 | 0.47 | 0.47 | 0.48 | 0.49
4 | 15 | G03 | 0.7 | 0.62 | 0.62 | 0.61 | 0.67 | 0.62 | 0.64 | 0.63 | 0.63 | 0.61 | 0.62 | 0.63 | 0.64 | 0.64 | 0.62
5 | 8 | G03 | 0.63 | 0.58 | 0.57 | 0.58 | 0.65 | 0.59 | 0.6 | 0.6 | 0.65 | 0.59 | 0.61 | 0.6 | 0.59 | 0.61 | 0.61
6 | 3 | G03 | 0.57 | 0.46 | 0.46 | 0.46 | 0.56 | 0.47 | 0.47 | 0.45 | 0.58 | 0.46 | 0.49 | 0.45 | 0.49 | 0.5 | 0.5
7 | 30 | G03 | 0.68 | 0.62 | 0.66 | 0.64 | 0.7 | 0.64 | 0.67 | 0.65 | 0.7 | 0.66 | 0.66 | 0.68 | 0.65 | 0.67 | 0.68
8 | 7 | G03 | 0.52 | 0.45 | 0.44 | 0.45 | 0.52 | 0.45 | 0.49 | 0.47 | 0.58 | 0.47 | 0.5 | 0.47 | 0.47 | 0.49 | 0.51
9 | 35 | G03 | 0.61 | 0.55 | 0.6 | 0.54 | 0.58 | 0.56 | 0.54 | 0.56 | 0.58 | 0.54 | 0.58 | 0.55 | 0.58 | 0.56 | 0.56
1 | 3 | G04 | 0.63 | 0.6 | 0.63 | 0.61 | 0.67 | 0.63 | 0.66 | 0.65 | 0.67 | 0.65 | 0.67 | 0.67 | 0.62 | 0.65 | 0.67
2 | 19 | G04 | 0.66 | 0.63 | 0.64 | 0.64 | 0.67 | 0.65 | 0.65 | 0.67 | 0.68 | 0.67 | 0.69 | 0.68 | 0.64 | 0.66 | 0.68
3 | 25 | G04 | 0.51 | 0.52 | 0.51 | 0.51 | 0.57 | 0.54 | 0.51 | 0.52 | 0.56 | 0.53 | 0.52 | 0.52 | 0.51 | 0.54 | 0.53
4 | 33 | G04 | 0.67 | 0.64 | 0.67 | 0.65 | 0.65 | 0.64 | 0.63 | 0.63 | 0.67 | 0.63 | 0.64 | 0.63 | 0.66 | 0.64 | 0.64
5 | 34 | G04 | 0.62 | 0.61 | 0.59 | 0.6 | 0.63 | 0.63 | 0.6 | 0.61 | 0.63 | 0.61 | 0.61 | 0.62 | 0.61 | 0.62 | 0.62
6 | 8 | G09 | 0.61 | 0.57 | 0.55 | 0.56 | 0.59 | 0.52 | 0.53 | 0.5 | 0.61 | 0.51 | 0.51 | 0.5 | 0.57 | 0.55 | 0.53
6 | 17 | G04 | 0.58 | 0.56 | 0.54 | 0.53 | 0.6 | 0.53 | 0.62 | 0.52 | 0.59 | 0.51 | 0.52 | 0.49 | 0.55 | 0.58 | 0.53
7 | 13 | G04 | 0.67 | 0.65 | 0.64 | 0.66 | 0.69 | 0.71 | 0.66 | 0.67 | 0.7 | 0.7 | 0.67 | 0.69 | 0.66 | 0.69 | 0.69
8 | 6 | G04 | 0.57 | 0.55 | 0.53 | 0.53 | 0.55 | 0.51 | 0.54 | 0.52 | 0.57 | 0.52 | 0.52 | 0.51 | 0.55 | 0.53 | 0.53
9 | 12 | G04 | 0.65 | 0.62 | 0.61 | 0.62 | 0.63 | 0.63 | 0.61 | 0.62 | 0.63 | 0.62 | 0.63 | 0.61 | 0.63 | 0.62 | 0.62
1 | 8 | G05 | 0.69 | 0.67 | 0.68 | 0.68 | 0.72 | 0.68 | 0.68 | 0.68 | 0.68 | 0.67 | 0.68 | 0.67 | 0.68 | 0.69 | 0.68
2 | 34 | G05 | 0.65 | 0.65 | 0.66 | 0.67 | 0.66 | 0.65 | 0.68 | 0.68 | 0.67 | 0.66 | 0.68 | 0.67 | 0.66 | 0.66 | 0.67
3 | 19 | G05 | 0.56 | 0.51 | 0.53 | 0.51 | 0.54 | 0.52 | 0.52 | 0.53 | 0.56 | 0.52 | 0.53 | 0.52 | 0.53 | 0.53 | 0.53
4 | 2 | G05 | 0.67 | 0.64 | 0.61 | 0.63 | 0.66 | 0.65 | 0.64 | 0.65 | 0.69 | 0.66 | 0.64 | 0.64 | 0.64 | 0.65 | 0.66
5 | 28 | G05 | 0.62 | 0.62 | 0.61 | 0.61 | 0.63 | 0.61 | 0.61 | 0.61 | 0.62 | 0.62 | 0.61 | 0.6 | 0.62 | 0.62 | 0.61
6 | 16 | G05 | 0.58 | 0.54 | 0.54 | 0.51 | 0.56 | 0.52 | 0.52 | 0.49 | 0.56 | 0.51 | 0.54 | 0.5 | 0.54 | 0.53 | 0.53
7 | 23 | G05 | 0.69 | 0.66 | 0.65 | 0.65 | 0.7 | 0.69 | 0.69 | 0.66 | 0.71 | 0.67 | 0.68 | 0.7 | 0.66 | 0.69 | 0.69
8 | 36 | G05 | 0.55 | 0.54 | 0.51 | 0.51 | 0.54 | 0.51 | 0.55 | 0.51 | 0.56 | 0.52 | 0.54 | 0.52 | 0.53 | 0.53 | 0.54
9 | 5 | G05 | 0.66 | 0.63 | 0.03 | 0.61 | 0.64 | 0.65 | 0.64 | 0.63 | 0.68 | 0.65 | 0.66 | 0.65 | 0.48 | 0.64 | 0.66
1 | 23 | G06 | 0.69 | 0.68 | 0.68 | 0.68 | 0.71 | 0.7 | 0.68 | 0.67 | 0.71 | 0.7 | 0.69 | 0.69 | 0.68 | 0.7 | 0.7
2 | 9 | G06 | 0.65 | 0.62 | 0.63 | 0.64 | 0.66 | 0.64 | 0.64 | 0.67 | 0.66 | 0.64 | 0.65 | 0.68 | 0.64 | 0.65 | 0.66
3 | 4 | G06 | 0.58 | 0.56 | 0.54 | 0.55 | 0.58 | 0.56 | 0.54 | 0.55 | 0.59 | 0.55 | 0.53 | 0.55 | 0.56 | 0.56 | 0.56
4 | 34 | G06 | 0.68 | 0.66 | 0.63 | 0.65 | 0.67 | 0.66 | 0.63 | 0.63 | 0.66 | 0.66 | 0.63 | 0.64 | 0.66 | 0.65 | 0.65
5 | 20 | G06 | 0.61 | 0.59 | 0.58 | 0.58 | 0.63 | 0.6 | 0.59 | 0.58 | 0.62 | 0.6 | 0.6 | 0.61 | 0.59 | 0.61 | 0.61
6 | 18 | G06 | 0.61 | 0.53 | 0.51 | 0.5 | 0.58 | 0.52 | 0.49 | 0.51 | 0.59 | 0.56 | 0.5 | 0.51 | 0.54 | 0.53 | 0.54
7 | 3 | G06 | 0.67 | 0.68 | 0.63 | 0.65 | 0.69 | 0.67 | 0.64 | 0.7 | 0.71 | 0.7 | 0.67 | 0.7 | 0.66 | 0.67 | 0.7
8 | 13 | G06 | 0.55 | 0.51 | 0.49 | 0.54 | 0.52 | 0.52 | 0.52 | 0.53 | 0.55 | 0.53 | 0.5 | 0.51 | 0.52 | 0.52 | 0.52
9 | 24 | G06 | 0.62 | 0.61 | 0.58 | 0.58 | 0.63 | 0.62 | 0.59 | 0.6 | 0.63 | 0.6 | 0.58 | 0.57 | 0.6 | 0.61 | 0.6
1 | 36 | G07 | 0.66 | 0.67 | 0.63 | 0.64 | 0.66 | 0.68 | 0.62 | 0.62 | 0.67 | 0.66 | 0.64 | 0.65 | 0.65 | 0.65 | 0.66
2 | 8 | G07 | 0.6 | 0.61 | 0.61 | 0.61 | 0.61 | 0.62 | 0.62 | 0.64 | 0.61 | 0.6 | 0.64 | 0.64 | 0.61 | 0.62 | 0.62
3 | 24 | G07 | 0.52 | 0.5 | 0.5 | 0.47 | 0.51 | 0.5 | 0.48 | 0.48 | 0.51 | 0.5 | 0.48 | 0.49 | 0.5 | 0.5 | 0.5
4 | 25 | G07 | 0.63 | 0.63 | 0.6 | 0.61 | 0.63 | 0.62 | 0.62 | 0.62 | 0.63 | 0.62 | 0.61 | 0.62 | 0.62 | 0.62 | 0.62
5 | 14 | G07 | 0.58 | 0.58 | 0.56 | 0.57 | 0.6 | 0.61 | 0.58 | 0.58 | 0.58 | 0.59 | 0.59 | 0.6 | 0.57 | 0.6 | 0.59
6 | 13 | G07 | 0.59 | 0.53 | 0.53 | 0.52 | 0.56 | 0.5 | 0.51 | 0.5 | 0.55 | 0.52 | 0.49 | 0.51 | 0.54 | 0.52 | 0.52
7 | 6 | G07 | 0.62 | 0.64 | 0.61 | 0.62 | 0.67 | 0.67 | 0.65 | 0.65 | 0.67 | 0.68 | 0.66 | 0.66 | 0.62 | 0.66 | 0.67
8 | 3 | G07 | 0.47 | 0.52 | 0.44 | 0.47 | 0.46 | 0.49 | 0.48 | 0.48 | 0.48 | 0.53 | 0.47 | 0.47 | 0.48 | 0.48 | 0.49
9 | 18 | G07 | 0.62 | 0.56 | 0.51 | 0.57 | 0.63 | 0.59 | 0.62 | 0.6 | 0.61 | 0.57 | 0.62 | 0.58 | 0.57 | 0.61 | 0.6
1 | 9 | G08 | 0.65 | 0.65 | 0.65 | 0.61 | 0.68 | 0.66 | 0.66 | 0.65 | 0.67 | 0.68 | 0.67 | 0.66 | 0.64 | 0.67 | 0.67
2 | 15 | G08 | 0.62 | 0.64 | 0.62 | 0.65 | 0.63 | 0.65 | 0.64 | 0.66 | 0.63 | 0.67 | 0.66 | 0.66 | 0.63 | 0.64 | 0.66
3 | 28 | G08 | 0.53 | 0.54 | 0.52 | 0.53 | 0.54 | 0.52 | 0.51 | 0.52 | 0.53 | 0.52 | 0.5 | 0.51 | 0.53 | 0.52 | 0.52
4 | 9 | G08 | 0.65 | 0.65 | 0.65 | 0.64 | 0.65 | 0.64 | 0.62 | 0.63 | 0.64 | 0.63 | 0.63 | 0.62 | 0.65 | 0.64 | 0.63
5 | 21 | G08 | 0.59 | 0.6 | 0.58 | 0.56 | 0.61 | 0.6 | 0.59 | 0.59 | 0.6 | 0.59 | 0.58 | 0.58 | 0.58 | 0.6 | 0.59
6 | 6 | G08 | 0.51 | 0.5 | 0.47 | 0.48 | 0.52 | 0.47 | 0.48 | 0.45 | 0.55 | 0.51 | 0.46 | 0.48 | 0.49 | 0.49 | 0.5
7 | 32 | G08 | 0.69 | 0.66 | 0.67 | 0.64 | 0.69 | 0.65 | 0.66 | 0.65 | 0.66 | 0.68 | 0.66 | 0.67 | 0.67 | 0.67 | 0.67
8 | 19 | G08 | 0.54 | 0.54 | 0.52 | 0.51 | 0.5 | 0.49 | 0.54 | 0.5 | 0.52 | 0.51 | 0.51 | 0.5 | 0.53 | 0.51 | 0.51
9 | 32 | G08 | 0.57 | 0.59 | 0.57 | 0.58 | 0.57 | 0.57 | 0.58 | 0.59 | 0.68 | 0.58 | 0.59 | 0.62 | 0.58 | 0.57 | 0.62
1 | 29 | G09 | 0.7 | 0.7 | 0.68 | 0.66 | 0.71 | 0.71 | 0.68 | 0.68 | 0.71 | 0.71 | 0.69 | 0.68 | 0.69 | 0.7 | 0.7
2 | 30 | G09 | 0.65 | 0.67 | 0.67 | 0.69 | 0.68 | 0.64 | 0.69 | 0.69 | 0.67 | 0.68 | 0.7 | 0.71 | 0.67 | 0.67 | 0.69
3 | 29 | G09 | 0.54 | 0.53 | 0.52 | 0.52 | 0.54 | 0.52 | 0.52 | 0.52 | 0.55 | 0.52 | 0.51 | 0.52 | 0.53 | 0.53 | 0.53
4 | 26 | G09 | 0.68 | 0.66 | 0.66 | 0.67 | 0.68 | 0.66 | 0.66 | 0.67 | 0.71 | 0.65 | 0.67 | 0.68 | 0.67 | 0.67 | 0.68
5 | 26 | G09 | 0.63 | 0.61 | 0.61 | 0.6 | 0.62 | 0.6 | 0.6 | 0.61 | 0.6 | 0.61 | 0.61 | 0.63 | 0.61 | 0.61 | 0.61
7 | 4 | G09 | 0.68 | 0.67 | 0.67 | 0.68 | 0.7 | 0.72 | 0.7 | 0.72 | 0.72 | 0.7 | 0.71 | 0.71 | 0.68 | 0.71 | 0.71
8 | 1 | G09 | 0.5 | 0.49 | 0.51 | 0.51 | 0.52 | 0.54 | 0.5 | 0.51 | 0.52 | 0.51 | 0.53 | 0.54 | 0.5 | 0.52 | 0.53
9 | 10 | G09 | 0.65 | 0.62 | 0.64 | 0.63 | 0.67 | 0.63 | 0.67 | 0.62 | 0.67 | 0.63 | 0.67 | 0.65 | 0.64 | 0.66 | 0.66
1 | 6 | G10 | 0.58 | 0.55 | 0.59 | 0.58 | 0.61 | 0.59 | 0.64 | 0.61 | 0.63 | 0.59 | 0.62 | 0.64 | 0.58 | 0.61 | 0.62
2 | 28 | G10 | 0.64 | 0.61 | 0.58 | 0.61 | 0.66 | 0.63 | 0.6 | 0.63 | 0.66 | 0.63 | 0.63 | 0.66 | 0.61 | 0.63 | 0.65
3 | 13 | G10 | 0.54 | 0.52 | 0.49 | 0.5 | 0.54 | 0.52 | 0.48 | 0.5 | 0.53 | 0.51 | 0.48 | 0.5 | 0.51 | 0.51 | 0.51
4 | 6 | G10 | 0.65 | 0.64 | 0.6 | 0.62 | 0.65 | 0.63 | 0.59 | 0.61 | 0.66 | 0.64 | 0.6 | 0.62 | 0.63 | 0.62 | 0.63
5 | 33 | G10 | 0.62 | 0.59 | 0.57 | 0.57 | 0.62 | 0.6 | 0.57 | 0.58 | 0.63 | 0.59 | 0.56 | 0.59 | 0.59 | 0.6 | 0.59
6 | 26 | G10 | 0.57 | 0.5 | 0.47 | 0.47 | 0.56 | 0.48 | 0.47 | 0.46 | 0.64 | 0.47 | 0.48 | 0.45 | 0.5 | 0.5 | 0.51
7 | 15 | G10 | 0.67 | 0.62 | 0.63 | 0.63 | 0.65 | 0.67 | 0.62 | 0.67 | 0.67 | 0.69 | 0.63 | 0.67 | 0.64 | 0.65 | 0.67
8 | 2 | G10 | 0.46 | 0.47 | 0.41 | 0.4 | 0.48 | 0.45 | 0.44 | 0.44 | 0.48 | 0.49 | 0.45 | 0.44 | 0.44 | 0.46 | 0.47
9 | 9 | G10 | 0.65 | 0.59 | 0.59 | 0.58 | 0.65 | 0.6 | 0.6 | 0.6 | 0.66 | 0.59 | 0.6 | 0.6 | 0.6 | 0.62 | 0.61
1 | 25 | G11 | 0.68 | 0.68 | 0.67 | 0.66 | 0.69 | 0.69 | 0.66 | 0.67 | 0.69 | 0.68 | 0.69 | 0.66 | 0.67 | 0.68 | 0.68
2 | 27 | G11 | 0.64 | 0.61 | 0.65 | 0.65 | 0.64 | 0.65 | 0.66 | 0.66 | 0.65 | 0.64 | 0.64 | 0.66 | 0.64 | 0.65 | 0.65
3 | 10 | G11 | 0.56 | 0.52 | 0.63 | 0.51 | 0.56 | 0.54 | 0.52 | 0.52 | 0.56 | 0.53 | 0.53 | 0.51 | 0.56 | 0.54 | 0.53
4 | 29 | G11 | 0.67 | 0.65 | 0.66 | 0.64 | 0.66 | 0.64 | 0.65 | 0.64 | 0.66 | 0.64 | 0.63 | 0.65 | 0.66 | 0.65 | 0.65
5 | 17 | G11 | 0.62 | 0.59 | 0.59 | 0.57 | 0.63 | 0.61 | 0.6 | 0.6 | 0.62 | 0.62 | 0.61 | 0.59 | 0.59 | 0.61 | 0.61
6 | 33 | G11 | 0.61 | 0.54 | 0.56 | 0.52 | 0.62 | 0.51 | 0.51 | 0.49 | 0.57 | 0.51 | 0.53 | 0.49 | 0.56 | 0.55 | 0.53
7 | 9 | G11 | 0.68 | 0.67 | 0.65 | 0.65 | 0.68 | 0.69 | 0.66 | 0.65 | 0.7 | 0.68 | 0.67 | 0.69 | 0.66 | 0.68 | 0.69
8 | 21 | G11 | 0.55 | 0.52 | 0.51 | 0.5 | 0.52 | 0.51 | 0.53 | 0.49 | 0.54 | 0.52 | 0.51 | 0.51 | 0.52 | 0.52 | 0.52
9 | 30 | G11 | 0.59 | 0.52 | 0.59 | 0.55 | 0.57 | 0.55 | 0.61 | 0.55 | 0.57 | 0.59 | 0.58 | 0.62 | 0.56 | 0.58 | 0.59
1 | 24 | G12 | 0.66 | 0.68 | 0.67 | 0.63 | 0.68 | 0.7 | 0.67 | 0.65 | 0.69 | 0.68 | 0.67 | 0.65 | 0.66 | 0.68 | 0.67
2 | 25 | G12 | 0.64 | 0.62 | 0.63 | 0.64 | 0.63 | 0.64 | 0.64 | 0.67 | 0.64 | 0.65 | 0.64 | 0.67 | 0.63 | 0.64 | 0.65
3 | 17 | G12 | 0.52 | 0.5 | 0.51 | 0.5 | 0.52 | 0.52 | 0.48 | 0.49 | 0.52 | 0.51 | 0.49 | 0.48 | 0.51 | 0.51 | 0.5
4 | 7 | G12 | 0.65 | 0.61 | 0.62 | 0.62 | 0.66 | 0.63 | 0.63 | 0.65 | 0.67 | 0.63 | 0.63 | 0.64 | 0.63 | 0.64 | 0.64
5 | 27 | G12 | 0.62 | 0.6 | 0.6 | 0.59 | 0.6 | 0.6 | 0.6 | 0.6 | 0.62 | 0.59 | 0.59 | 0.6 | 0.6 | 0.6 | 0.6
6 | 34 | G12 | 0.56 | 0.54 | 0.52 | 0.5 | 0.56 | 0.49 | 0.51 | 0.5 | 0.55 | 0.49 | 0.52 | 0.49 | 0.53 | 0.52 | 0.51
7 | 18 | G12 | 0.65 | 0.65 | 0.65 | 0.66 | 0.66 | 0.66 | 0.65 | 0.67 | 0.68 | 0.64 | 0.66 | 0.65 | 0.65 | 0.66 | 0.66
8 | 5 | G12 | 0.49 | 0.47 | 0.45 | 0.46 | 0.47 | 0.46 | 0.5 | 0.46 | 0.51 | 0.47 | 0.48 | 0.47 | 0.47 | 0.48 | 0.48
9 | 16 | G12 | 0.62 | 0.6 | 0.59 | 0.59 | 0.61 | 0.58 | 0.59 | 0.61 | 0.62 | 0.62 | 0.6 | 0.6 | 0.6 | 0.59 | 0.61